Question and Answer on Trend Report

Introduction

This document aims to explain and clarify some questions related to Trend Report that are outlined in Article 88 of the Regulation (EU) 2017/745 on medical devices (MDR) and Article 83 of Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR). A common understanding and clarification related to Trend Report is deemed essential for an effective and harmonised implementation of the vigilance requirements under the MDR and IVDR.  To that end, this document is intended as a supporting tool for use by competent authorities, economic operators and other relevant parties to be read, in conjunction with the MDR and IVDR, to establish a common understanding of the MDR and IVDR vigilance requirements and obligations.

The requirements on Trend Report, as defined in the Article 88 of the MDR, are applicable from the 26th May 2021 and this document is already an effective guidance.

The requirements on Trend report, as defined in the Article 83 of IVDR, are applicable from 26th May 2022. The Trend Report form will be available when Eudamed becomes fully functional. The current Trend report form should be used until that.

The Trend Report under the MDR requires a more consistent, standardized and systematic review of all incidents and expected side effects apart from serious incidents, reported by medical device manufacturers. As far as IVDR requirements are concerned, a standardized and systematic review of all incidents, apart from serious incidents, and expected erroneous results is needed.

This document takes into account the “preparatory document for the PMSV WG meeting on the implementation of Article 87(1) MDR” which highlights that it is not necessary to distinguish between ‘old’, ‘legacy’ and ‘MDR’ devices for the purpose of defining the vigilance reporting obligations.

The Terms and Concepts, useful for the Trend Report, as outlined in Section 2 of Chapter VII of the Regulation (EU) 2017/745 on medical devices, are defined on the document Questions and Answers on Vigilance Terms and Concepts (add reference when available).

The concepts defined in this document are applicable to the corresponding articles of the Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR), when IVDR-specific Q&A are not included.

- What shall a manufacturer report with a Trend Report?

The Medical Device Regulation 2017/745 Article 88 (1) requires manufacturers to report as Trend Report, to the Evaluating Competent Authority, any statistically significant increase in the frequency or severity of incidents, that are not serious incidents, or that are expected undesirable side-effects that could have a significant impact on the benefit-risk analysis.   The statistically significant increase shall be established in comparison to the foreseeable frequency or severity of such incidents in respect of the device, or category or group of devices, in question during a specific period as specified in the technical documentation and product information.

The In vitro Medical Device Regulation 2017/746 Article 83 (1) requires manufacturers to report as Trend Report, to the Evaluating Competent Authority, any statistically significant increase in the frequency or severity of incidents or expected erroneous result established in comparison to the stated performance of the device as referred to in points (a) and (b) of Section 9.1 of Annex I and specified in the technical documentation and product information.

The Competent Authority shall be notified  in cases where a statistically significant increase  of incidents or expected undesirable side effects (MDR) or expected erroneous result (IVDR) is detected and that change could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. Trends should be identified by the manufacturer as they can be indicative for a change in the risk-benefit ratio.

- When a manufacturer shall report an incident, an expected undesirable side effect or an expected erroneous result to a Competent Authority using a Trend Report?

The manufacturer has to record incidents and expected undesirable side effects in accordance with Articles 83-86 of the MDR prescription for the post market surveillance.

As specified in the Article 83 (2) the manufacturer in its post market surveillance system has to gather, record, analyze, actively and systematically, any relevant data on the quality, performance and safety of a device throughout its entire lifetime, in order to be able to draw the necessary conclusions and to determining, implementing and monitoring any preventive and corrective actions. As also specified in the Article 83 (3), data gathered by the manufacturer's post-market surveillance system shall, among other things, be used to detect and report trends in accordance with Article 88.

Whenever a statistically significant increase in the frequency or severity of not serious incidents or of expected undesirable side effects or expected erroneous result is detected and results in a change of the risk evaluation and have a significant impact on the benefit-risk analysis, the manufacturer shall draw up a trend report as defined in the Article 88/83 of the MDR/IVDR.

A trend report can be drawn up when the manufacturer has performed a monitoring of incidents over time and as compared the gravity and the frequency of the occurrences with the threshold defined during the design process and documented in the risk analysis.

- What happen if a NEW () undesirable unexpected side effect or incident is identified?

Undesirable unexpected side-effects or incident are events that were not considered or addressed by the manufacturer in its risk analysis should be reported in accordance with Article 87 of the MDR (serious incidents) and  documented and analyzed during the update  the risk analysis and the risk management file.

The update of the risk-benefit assessment should take the identified new risks in consideration. Any residual risks should be communicated to the users/patients via the labeling of the product.

Whenever a statistically significant increase in the frequency or severity of incidents or of /unexpected undesirable side effects is identified and confirmed according the manufacturer internal procedures, the manufacturer should conduct a full assessment of the trend on   impact of the trend  on the benefit-risk ration of the product and shall draw up a trend report as defined in the Article 88/83 of the MDR/IVDR.

- How to link the PMS plan with the trend reporting processi?

As per Article 88 (1) (MDR), the manufacturer shall specify, in the post-market surveillance plan referred to in Article 84 for MDR:

-  how to manage the incidents subject to the trend report

-  the methodology used for determining any statistically significant increase in the frequency or severity of such incidents

- the observation period.

Article 84 (MDR) specifies that the requirements for the post-market surveillance plan are set out in Section 1 of Annex III which says that “methods and protocols to manage the incidents subject to the trend report as provided for in Article 88, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period.”

As per Article 83 (1) IVDR the manufacturer shall specify, in the post-market surveillance plan referred to in Article 79 of the IVDR:

-  how to manage the incidents subject to the trend report

- the methodology used for determining any statistically significant increase in the frequency or severity of such events or change in performance

- the observation period.

- How to link the trend reporting process to general Quality Management System?

In line with Article 88 of MDR (Article 83 IVDR) on trend reporting, main elements of the Quality Management System (QMS) already in place (Complaint Handling, Vigilance System, Trend Review, Quality Improvement Review, Risk Management Process, Clinical Evaluation Process, and Corrective Action and Preventative Action - CAPA system) are to be used for the implementation of the trend reporting requirement and the latter can be used to reinforce the Quality Management System.

The risk management and clinical evaluation processes under which manufacturers operate are essential key elements in trend reporting being the basis of determining:

- when trends become significant,

- when the acceptable (lowest) risk level (severity-harm combination) might become threatened,

- when the risk-benefit analysis of residual risk has led or may lead to unacceptable risks to health or safety of patients,

- when the residual risk does not constitute an acceptable risk when weighed against the benefits to patient,

- when actions are necessary by manufacturers to address unacceptable risks.

- Which incidents/undesirable side-effects/ expected erroneous result from a geographic perspective have to be included in Trend Reporting? ?

The manufacturer has to record incidents and expected undesirable side effects in accordance with Articles 83-86 of the MDR prescription for the post market surveillance.

In order to identify a trend, which is based on a statistical data, the manufacturer has to take into account the worldwide incident.

Therefore, the detection of trend is based on the data collected about the incidents related to the device occurred in Europe and world wild. The manufacturer should prepare a trend report when at least one incident occurs in the EEA+TR+IX.

- Does a trend report refer only to one device or also to a group of devices?

As indicated in the Article 88 (1) of the MDR, a trend report could be performed as for a single device as for a category or group of devices.

For this reason, in the Section 2 of the Trend Report Form the manufacturer can select the device scope type choosing between the medical device high level EMDN code i.e. group of medical devices, Basic UDI-DI/Eudamed DI, UDI-DI/Eudamed ID and UDI-PI/Eudamed ID (lot/batch number). The old devices could be reported only by the level of device name.

Consider that only one choice is possible inside the device scope type.

- Could the device scope and the scope of the trend be changed during the trend report process and the trend report life cycle?

When a trend report is performed and a device scope is defined it has to be kept the same during the trend report process and trend report life cycle. This is because a trend report is based on statistical data collected by the manufacturer with a specific methodology and they have to be referred to a specific device scope in order to avoid any confusion and to allow to the Evaluating Competent Authority to carry out the appropriate assessments on the data provided.

Indeed, the device scope has to remain unchanged during the trend life cycle and updating the device scope is not possible. If a new device scope is identified for the same device/category/group of devices, it is necessary to create a new trend report with its own process and life cycle.

It is firstly important to clarify that the Trend report should be based on one specific issue (trend report scope) identified by the manufacturer, based on a benefit/risk evaluation related to that specific problem, and which refers to one device or to a group/family/category of device (device scope). Only if the assessment, under which the trend report is based on, remain unchanged. (For example: a new batch of device is identified).

- If a device has different problems, how should they be reported?

If a device has different problems, each problem should be reported in a single trend report.

- With reference to Article 88/83 of the MDR/IVDR and the criteria for trend reporting, what is meant by a ‘significant impact on the benefit-risk analysis’?

A ‘significant impact on the benefit-risk analysis’ is whenever the increase in the frequency or severity of incidents, that are not categorized as serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation (EN ISO 14971). A significant increase of the severity could be measured as possible consequence of the incident and / or increase of the probability of occurrence of harm in such a way that could lead to unacceptable Residual Risk and significant impact on the benefit-risk analysis.

In line with Article 88/83 of MDR/IVDR and in reference to Articles 1 and 5 of Annex I / Chapter I, the final decision to report is complemented by health hazard assessment in which the residual risk is weighed against the intended benefits.

- What is the “threshold” mentioned in the section 3 (b)?

The established thresholds refer to the threshold identified by the manufacturer in its risk assessment. In this section the manufacturer should describe how he has figured potential risk and which signal he detected.

More information in Annex I of this document.

- Which are the documents related to the Trend Report?

When the manufacturer issues a Trend Report, they have to take into account 2 documents:

-Eudamed Trend Report Form. This document submitted directly in Eudamed when fully functional identifies: the Member States where the incidents occurred, the administrative information, description of the devices and the detected trend (background, established threshold, methodology, observation period, root cause analysis, actions to be taken or already implemented to the reduce the risks for the users/patients ….).

- Trend Report document. Additional document  should be added as required. This document is a file attached to the Eudamed Trend Report Form and includes more detailed information about the affected devices, the method and description of the detected trend and its investigations and conclusions.

As mentioned above, the Eudamed Trend Report Form (and its related helptext) will be applicable only when Eudamed becomes fully functional. In the meantime, although the MDR Trend report requirements are applicable, alternative administrative technical solution have to be adopted as per MDCG 2021-1 Guidance on harmonised administrative practices and alternative technical solutions until EUDAMED is fully functional.

- How the manufacturer shall notify the Trend Report to an Evaluating Competent Authority?

The manufacturer shall notify the trend report using the Eudamed Trend Report Form in the Eudamed system as specified in the Article 88 (1) of the MDR and Article 83 (1) of the IVDR.

As per art. 92 (6) MDR and art. 87 (6) IVDR the trend reports shall be automatically transmitted upon receipt via the electronic system to the competent authorities of the Member States in which the incidents occurred.

Initially the manufacturer provides an initial (or a Combined Initial and final) Trend Report Form. The manufacturer provides Trend report follow up and Final trend report together with the trend report document in attachment

Until Eudamed is fully functional Manufacturers must submit trend reports to the respective / relevant national vigilance data base. The current Trend report form should be used until  Eudamed is fully functional.

- Can a Trend Report lead to the definition of a corrective action/preventive action?

During the assessment of the data basis of the trend report, the manufacturer can  issue a corrective action including field safety corrective action/preventive action to prevent or reduce risk for patient or user of the device. In this context the manufacturer should notify a Final trend report and initiate an appropriate action (i.g. FSCA, CAPA, update of labeling, training).

If the evaluating CA is not satisfied with the action taken, they shall require additional corrective actions as appropriate as specified in article 88 (2)/ 83 (2) of MDR/IVDR.

- How could the manufacturer manage the Custom Made Device CMD in the scope of Trend Report?

The same process related to the identification of a Trend applies also on the CMDs.

With the absence of stated exceptions, CMD manufacturers must meet nearly all of the MDR requirements.

As defined also in the guidance MDCG 2021-3, for risk management, post-market surveillance and clinical evaluation life cycle processes as defined by the MDR, CMD manufacturers should apply these obligations to groups of devices with the same intended purpose, materials used, process utilized, same principal design etc. and not to each individual CMD.

As a consequence, also the requirements of Article 88 of the MDR could be applied.

- What are some examples for reporting a Trend Report?

In order to further clarify what shall be reported in a trend report, some examples are provided below:

- A trend of expected undesirable side effects causes a risk zone changes from low to medium, or an increase from medium to high, within the risk matrix as suggested in accordance with ISO 14971, Medical devices, Application of risk management to medical devices;

- A number of similar incidents, according Art. 2 (64), that are not serious incidents reported by different healthcare facilities shows a statistical increase with a daily or monthly frequency;

- …… add other examples also for IVD

Annex I

Threshold

A ‘significant impact on the benefit-risk analysis’ is whenever the increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation. An example of this can be, a risk zone changes from low to medium, or and increase that is from medium to high, within the risk matrix as suggested in accordance with ISO 14971:2019.

Trigger levels of risk for trend reporting should be defined at the beginning of the designing and manufacturing process in accordance with the risk-benefit analysis and reference to them should be reported in trend report.

In particular,

Risk can be considered as the results of two components:

-  probability of damage occurring;

- consequences of the damage, meaning gravity of the damage.

Benefit is defined as the benefit of a medical device is related to the likelihood and extent of the health improvement expected from its use.

Risk can be considered as the results of three components:

- the expected performance during its clinical use;

- the clinical outcome expected from this service;

- factors relevant to the risks and benefits of other treatment options.

As the standard ISO 14791 do not define exactly the “acceptable risk” the evaluation is demanded to the manufacturer who should have a suitable system in place for proactive scrutiny.

Some qualitative methods can be identified in agreement with the EN ISO 14971.

Example of matrix for qualitative risk assessment according to EN ISO 14971 Medical devices, Application of risk management to medical devices.

|  | Negligible | Moderate | Significant |
| --- | --- | --- | --- |
| High |  |  |  |
| Medium |  |  |  |
| Low |  |  |  |

Qualitative seriousness levels

Qualitative probability levels

Defining “qualitative seriousness levels” as:

- significant: death or loss of function or structure

- moderate: reversible or minor injury

- negligible: produces no injury or causes minor injury

and “qualitative probability levels”:

- High: it tends to occur often, very frequently

- Medium: it can occur but infrequently

- Low: rare, remote

Acceptable risk

Not acceptable risk

Example of matrix for semi-qualitative risk assessment according to ISO 14971, Medical devices, Application of risk management to medical devices.

Semi- Qualitative probability levels

Qualitative seriousness levels

|  | Negligible | Minor | Serious | Critical | Catastrophic |
| --- | --- | --- | --- | --- | --- |
| Frequent |  |  |  |  |  |
| Probable |  |  |  |  |  |
| Occasional |  |  |  |  |  |
| Remote |  |  |  |  |  |
| Unlikely |  |  |  |  |  |

Defining “qualitative seriousness levels” as:

- Catastrophic: causing the death

- Critical: produces permanent impairment or a life-threatening injury

- Serious: produces permanent impairment or injury that requires professional medical intervention

- Minor: produces temporary impairment or injury that does not require professional medical intervention

- Negligible: temporary discomfort

A trend report should be initiated when the events, as classified by the manufacturer, impose changing the level identified, moving from “white zone” to “red zone”, in terms of frequency or in term of gravity, or a combination of the two.

Until Eudamed becomes fully functional national systems should be used as defined in the MDCG 2021-1 Rev.1 Guidance on harmonised administrative practices and alternative technical solutions until EUDAMED is fully functional

A distinction between ‘legacy’ and ‘old’ devices may remain relevant for the registration of the devices in Eudamed, i.e. how to link a serious incident report registered in the vigilance module to the registration of a device in the UDI / device module.  This issue has been addressed by MDCG guidance 2019-5.

referred to in Sections 1 and 5 of Annex I MDR and IVDR

EN ISO 14971 Medical devices, Application of risk management to medical devices.

Probability of occurrence of harm = the probability of hazard or hazardous situation multiplied by the probability of hazard or hazardous situation causing harm.  (EN ISO 14971:2019 definition)

Residual Risk = combination of the probability of occurrence of harm and the severity of that harm. (EN ISO 14971:2019 definition).

Refer also to MDCG 2021 – 3 guidance “Questions and Answers on Custom-Made Devices” Question 8 and 9.

Question 8 MDCG 2021 -3: 8. What are the obligations of CMD manufacturers according to the MDR?

With the absence of stated exceptions, CMD manufacturers must meet nearly all of the MDR requirements. Although the authorised person issuing the written prescription primarily determines the design and intended purpose of the CMD, it is the responsibility of the CMD manufacturer to consider which MDR Annex I requirements are applicable to the device at hand.

Additionally, and as any other medical device manufacturer, CMD manufacturers shall establish, document, implement, maintain, keep up to date and continually improve a quality management system (QMS) that shall ensure compliance with the MDR in the most effective manner and in a manner that is proportionate to the risk class and the type of device. The QMS must address all elements described in Article 10(9) of the MDR.

To implement an MDR compliant post-market surveillance system, the CMD manufacturer should establish appropriate communication channels with relevant healthcare providers/professionals or patients to receive feedback on the quality, performance and in particular the clinical performance and safety of the devices in the field.

For risk management, post-market surveillance and clinical evaluation life cycle processes as defined by the MDR, CMD manufacturers should apply these obligations to groups of devices with the same intended purpose, materials used, process utilised, same principal design etc. and not to each individual CMD.

In accordance with Article 87(1) of the MDR, CMD manufacturers shall report to the competent authorities any serious incidents and/or field safety corrective actions as soon as they learn of them.