| NOMA |  | 544-550 | 16 |  | Te |  | We do not agree with the proposed text regarding the evaluating CA for FSCAs. | We suggest that the evaluating CA for FSCAs is the CA of the MS in which the manufacturer or authorized representative is located, and that the FSN draft is sent exclusively to that evaluating CA. According to article 89(8) the intention is that the content of the FSN shall be consistent in all MS. To circulate the FSN draft to CAs of all the MS in which the device is distributed, collect responses, and edit the FSN accordingly, will delay the manufacturer’s process of distributing the FSN. | Accept. See revised text |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FAMHP |  | 57 | 1. |  | TE |  | Figure While there is not one process to handle complaints, from our experience most manufacturers use a slightly different one: Complaint? -> Incident? -> serious incident? They first device if it is an incident (and as such need a different treatment in regulatory affairs/PMS/… department and afterwards decide if it is a serious incident. That is also logical as the difference between complaint and incident is on manufacturer malfunction of the device + side effect. While the difference between incident and serious incident is about the unusual/unexpected part of the event and the clinical consequences. | Add a box between complaint and serious incident with a “Does the complaint fulfil the criteria of an incident” Potential questions are: - Did the device perform correctly as intended? - What is the relationship between incident and MF device And a box “Does the incident fulfil the criteria of a serious incident” - Is it an expected, documented side-effect - what was the outcome or potential outcome of the incident and add the YES/NO arrows after this box | Accept. See revised flowchart |
| FAMHP |  | 112-114 | 2. |  | TE |  | “side effects that are not quantified in the technical documentation and/or not addressed in the information supplied by the manufacturer to the end user” This does not define an incident, this defines a serious incident already. According to 2 (64) definition on incident it should be “any undesirable side-effect.” Potential corrections are clarifying that this bullet point is already on serious incidents or adding the text to another point (in which C would be the most logic, after the possible casual relationship between the side-effect and the device has been established it can be checked if this expected and well documented. | See comments for 2 possible suggestions, in this only the second is elaborated. Change bullet point in A. to: - any undesirable side-effect” Add to C. In case of an undesirable side-effect, the side-effect is not already quantified in the technical documentation and/or addressed in the information supplied by the manufacturer to the end user, i.e. the instructions for use, labelling or other accompanying documentation | Partially accept. Bullet point in A is corrected. Addition to C is not included in the text, however see revised flowchart, which includes an elaboration of side- effects and revised text in question 7 on side –effects. |
| FAMHP |  | 129-138 | 2. |  | GE |  | Shouldn’t the foetal distress, foetal death or a congenital physical or mental impairment or birth defect; be mentioned somewhere? These are in fact serious deteriorations in the state of health (or death) only not of the treated patient. A footnote to clarify that person with incidents/serious incidents is not limited to the patient but includes fetus, users (i.e. doctor/nurse) or any other person. Could be interesting. | Discretion of the taskforce | Foetal distress added in the text. On the proposal to include clarification in footnote, this cannot be added as this may open up for a broader discussions such as definition of a fetus and whether a fetus qualify as a third person. This point was not discussed with the TF, but internally in the COM. |
| FAMHP |  | 176 | 3. |  | GE |  | Examples of deterioration: - sensor drift - UV degradation - Elasticity changes (increase or decrease) - Wear | / | Accept |
| FAMHP |  | 439 | 13 |  | GE |  | Shouldn’t first the reason this field exist be explained prior to its usage? It is also not only additional information that can change this. It is possible that due to the strict timelines a manufacturer had only time to do a superficial analysis, and needs to report | It is possible that a manufacturer has not established within 2, 10, 15 days of the awareness of a potential serious incident if it is reportable or not. According to article 87 (7), the manufacturer is obliged in case of uncertainty to report such a potential serious incident. Further investigation after the 2, 10, 15 deadline could clarify that the event did not meet the serious incident criteria. Similar the manufacturer’s analysis of additional information, received after the submission of an initial MIR, can also conclude that the event is not a serious incident. In those cases the manufacturer can close these reported events by selecting the box ‘non-reportable’ in section 1.2 (e) of the MIR form. However, if the root cause has not been clearly established or identified this does not by itself qualify the case as non-reportable. …” | Accept see revised text. |
| FAMHP |  | 495 | 14 |  | ED |  | “It should be noted that, a corrective action inkling advice or instructions provided by a manufacturer can…” The usage of inkling is a bit weird in this sentence as inkling means “a feeling that something is true or likely to happen, although you are not certain:” Could be wrong as not a native speaker | Change inkling with accompanying “It should be noted that, a corrective action accompanying advice or instructions, provided by a manufacturer can…” | See revised text |
| FAMHP |  | 529-536 | 16 |  | TE |  | This should include the Member State in which the manufacturer has its registered place of business (see article 92.7) | “the competent authorities of the Member State(s) in which concerned devices are distributed or patients are located. Additionally the competent authority in which the manufacturer or his authorized representative has its registered place of business should be informed if not already among the Member States with concerned devices or patients. “ Or another formulation. | Accept. Previous text removed to avoid any misunderstandings. |
| FAMHP |  | 652 | Annex |  | GE/ED |  | We don’t 100% understand this table. Our issue is mainly with the first box below the purple arrow. The first point seems to be closer to criteria C; the last can be on all 3 criteria. We are also missing side-effects but it could be that it is in the box, again we don’t 100% understand the box. From an editorial point it could be interesting to make all boxes on the same level the same size (in height) + aligning them. | / | The annex is deleted. |
| MPA | 25 |  | Q1 |  | Ed | Consider to revise, for reason of clarity, include “to competent authorities” |  | Not all incidents are reportable to competent authorities, | Accept |
| MPA | 28 |  | Q1 |  | ed | Consider to revise, for reason of clarity, include reference to MDR article 87” |  | In accordance with Article 87 | Accept |
| MPA | 29 |  | Q1 |  |  | Include reference to MDR article 83 |  | … post-market surveillance (Article 83) and be reported in… | Accept |
| MPA | 43 |  | Q1 |  | Ed | Replace should with shall as this is stated in MDR reference 87(7) |  | Shall | The word ‘must’ is used when the text refers to MDR requirements. Whilst the previous version of the document did, use the word ‘shall’, MDCG guidance documents cannot use the word shall, unless the text is a direct quotation of the MDR. |
| MPA | Box 1 |  | Figure 1 |  | Ed | Information about incidents and serious incidents are not from received from competent authority, but may be received via a competent authority |  | Via a competent authority | Accept |
| MPA | 70 |  | Figure 1 |  | te | Difference between should provide a MIR for the serious incident and SHALL be reported in accordance with MDR. |  | The manufacturer shall report the serious incident in accordance with Article 87 (1) of the MDR within the time requirements outlined in Article 87 (4)(5)(3) of the MDR. The incident should reported in a Manufacturer Incident Report (MIR). | Accept. The word ‘must’ in added to indicate MDR requirement of reporting serious incident and see FN 8 for reference to the MIR report. |
| MPA | 71 Box under No Incident 2(64) |  | Figure 1 |  | te | Please emphasize that for incidents, trend reporting is a possible requirement (hence, shall). |  | The manufacturer does not have to provide a Manufacturer Incident Report (MIR). For incidents that are not serious incidents, the manufacturer shall comply with the requirements for trend reporting in accordance with Article 88 of the MDR. | Accept. See revised text. |
| MPA | 135 |  | Q2 |  | te | In some cases, an incident has been argued not to be reportable as it did not lead to hospitalization, although medical intervention was required. National healthcare systems seem to work differently, and these differences should not be what decides if an incident is reportable or not. |  | Replace “a condition necessitating hospitalisation” with “a condition necessitating professional medical care” An alternative is to complement the current text with “or a condition necessitating professional medical care” | Accept. Proposal added. |
| MPA | 149 |  | Q2 |  | te | Examples of a serious public health threat should also include non-contagious examples |  | Add following examples as bullets: toxic compounds with a potential negative effect on humans, non-sterile devices labled as sterile, cyber attacks, | Accept |
| MPA | 162 |  | Q2 |  | ed | Difficult to understand (err), please consider to rephrase |  | In complex situations, it should be assumed that the device may have or potentially could have contributed to the serious incident. In case of doubt on the reporting requirements, there should as a general principle always be a pre-disposition to report rather than to not report. | See revised text. |
| MPA |  |  | Annex 1 |  | ed |  |  | Replace ”compliant” by ”complaint” thorough Annex 1 | Annex deleted |
| MPA |  |  | Annex 1 |  |  | Orange box, it is a requirement, hence “shall” and not should |  | Shall be reported to the relevant… | Annex deleted |
| MPA |  |  | Annex 1 |  |  | Remove ”it is a criteria that” |  | An incident (Article 2(64)) has occurred | Annex deleted |
| MPA |  |  | Annex 1 |  |  | Make this a figure in Q2 and consider to make the following adjustments for reasons of clarity: |  | Make the first blue box a headline. Remove the blue arrows. Move the lower text boxes above the blue (and orange) “Criteria”-boxes, and place arrows from the text boxes to the “criteria”-boxes. | Annex deleted |
|  |  |  |  |  |  | For additional clarity, and learning purposes, consider to add information of what is expected if the A, B or C criteria is not fulfilled. See draft suggestion as attached file in the email |  |  | The information provided in the annex has been incorporated in the text (see question 2 – Criterion C). The additional information related to what is expected if the criteria (A-B) are not fulfilled, can also be included in the text if deemed necessary by SE? |
| MPA | 173 |  | Q3 |  | Ed | Grammar, “hence function in order to achieve…” should state that the device does not function as expected to achieve its intended purpose… |  | Is unable to uphold its performance (Article 2(22) MDR) and hence fail to achieve its intended purpose as stated by the manufacturer. | Accept. Text revised |
| MPA | 176 |  | Q3 |  | te | Include deterioration |  | Malfunctions or deteriorations | Accept |
| MPA | 189 -191 |  | Q3 |  | Te | For sake of clarity, please consider to rephrase in line with this suggestion. |  | A device which is unable to function or achieve its intended purpose when used in accordance with the manufacturer´s instructions for use should always be investigated by the manufacturer. Such events can cause or lead to serious incidents and be reportable in accordance with Article 87 or in the case of incidents (2(64) in accordance with article 88 MDR. | Accept. See revised text |
| MPA | 192 - 195 |  | Q3 |  | te | Please revise this section, it adds confusion, and it is not correct according to iso 14971 and iso 62366. Avoid the use of “misuse” as the term is part of “reasonably foreseeable misuse” and not to be used alone. What is the aim of this section, row 192-195 and what is the reference for this interpretation? Do you mean that for the purpose of Article 2(64), abnormal use is excluded as cause for device malfunction or deterioration? MPA regard device malfunction or deterioration in accordance with article 2(64) to include normal use (correct use and use errors). |  |  | Text changed. |
| MPA |  |  | Q5 |  |  | This section was not correct according to iso 14971 and en 62366:2008. Please find our suggestion |  | A ‘use error’ is when the user’s action or lack of such, while using the device, leads to a different result than that expected by the user or intended by the manufacturer. Use errors can be a result of a mismatch between different factors occurring during normal use of a device such as; the characteristics or profile of the user, the user interface, or the nature of the task and/or the environment or circumstances in which the device is used. Use errors can be categorised as being caused by attentional failure (slips), memory failures (lapses) or rule-, knowledge-based and nescient errors (mistakes). Use errors are included in normal use. Abnormal use is the deliberate violation of the intended use of a medical device. Examples of abnormal use can be: - off-label use of the device, e.g. a doctor which based on a medical assessment, decides to use a device for a different indication or in a different manner than indicated in the accompanying device documentation to suit the needs of a specific patient. - an intentional disregard of the information for safety, provided in for example the instructions for use, product labelling or other accompanying documentation with relevant device information; - reckless use. Use errors which have led, might have led/might lead to the outcome of a serious incident must be reported in accordance with Article 87(1) of the MDR. Given the nature of a use error, the user might not be aware that an error has occurred or that the device is not working as intended. Information to manufacturers that are related to use errors (independent of whether they are related to the ergonomic features of a device) or abnormal use, should be recorded and handled within the manufacturer’s quality and risk management system and included in the technical documentation on post-market surveillance in accordance with Annex III to the MDR. | Partially accept. See revised text. Use errors that are not caused by ergonomic features do not fall within the scope of an incident and are not reportable under Article 87(1) MDR . |
| MPA | Figure 2 |  | Q5 |  |  | The current figure does not add additional clarity. Consider to adjust or replace this figure, see suggestions in the attached document (figure adjusted from EN 62366:2008, figure B1, page 32) |  |  | Accept. Delete figure |
| MPA | 295 |  | Q7 |  |  | Include ”might have been caused” |  | Based on the health effect caused or that might have been caused by… | Accept |
| MPA | 310 |  | Q8 |  |  | Remove ”practicable” from “eliminate or reduce the risk as low as possible and practicable” as this is not used in MDR, see chapter one of MDR Annex I |  | eliminate or reduce the risk as low as possible | Accept |
| MPA | 322 |  | Q10 |  | Ed | For consistency, in the headline, include a ´ character in front of the question mark |  | ´without undue delay´? | Accept |
| MPA | 351 |  | Q11 |  | Te | Remove the row 351 to 356 as this contradicts the example on row 385-387 and Question 12. The reporting period begins the day after the information was received (see awareness date Q12) and ends on the… |  |  | See revised text. |
| MPA | 348-358 |  | Q11 |  | Te | Adjust as required when row 351 – 356 is removed. |  | The reporting period begins the day after the information was received (see elaboration on awareness date in question 12) and ends on the 2nd, 10th, or 15th day thereafter. | See revised text |
| MPA | 413 |  | Q12 |  |  | Grammar, instead of “to be reportable”, use “is reportable” or “is to be reported” |  | “is reportable.” or “is to be reported.” | Accept. Changed to ‘is reportable’ |
| MPA | 458 |  | Q14 |  | Ed | Grammar: “an”, not “a” FSCA |  | An FSCA | Accept |
| MPA | 487 |  | Q14 |  | Ed | Grammar, avoid -ing form |  | Advice related to changes… | Accept |
| MPA | 495 |  | Q14 |  | Ed | Spelling: inkling |  | including | Accept |
| MPA | 495 - 499 |  | Q14 |  | te | The use of the term corrective action (2(67)) for the risk mitigation measures implemented within the FSCA should preferably be exchanged. An FSCA is a correction and its mitigation measures i.e., recall of devices from the market, does not necessarily eliminate the cause (root causes and contributing factors) of a potential or actual non-conformity (or other undesirable situation) Replace “corrective action” with “FSCA”, or “measure” as applicable. |  |  | Accept |
| MPA | 495 |  | Q14 |  |  | As reasoned above, see adjustment of the text: |  | It should be noted that, advice or instructions provided by a manufacturer in the FSN can include modifications… | Accept |
| MPA | 499 |  | Q14 |  |  | As reasoned above (Q14 row 495-499), see adjustment of the text: |  | clinically unjustifiable to explant the device. Therefore, special patient follow-up or treatment, irrespective of whether any affected un-implanted devices remain available for return, constitutes measures to be included in an FSCA. | Accept |
| MPA | 533 |  | Q16 |  |  | See our reasoning on Q14 (row 495-499) about the use of the term corrective action (2(67)). Corrective action is not the same as FSCA, in this context FSCA should be used. |  | competent authorities in Member States affected by the field safety corrective action | Accept |
| MPA | 538 - 539 |  | Q16 |  |  | Evaluating CA is not the same as coordinating CA. Use “shall” inform, as this is required in accordance with 89(9). |  | The coordinating competent authority shall inform the manufacturer… | Accept. |
| MPA | 556 |  | Q17 |  |  | Include reference to the Regulation (EU) 2021/2078 where information can be found about how to act in case of technical unavailability or malfunction of Eudamed – after Eudamed has been “released” |  | In the event of technical unavailability or malfunction of Eudamed, information on how to act can be found in “COMMISSION IMPLEMENTING REGULATION (EU) 2021/2078 of 26 November 2021 laying down rules for the application of Regulation (EU) 2017/745 of the European Parliament and of the Council as regards the European Database on Medical Devices (Eudamed). | Accept |
| MPA | 558 |  | Q18 |  |  | Is this question relevant for this document as the aim of the article relates to the assessments performed by CA. |  | Remove the question. | Can be removed if TF agrees. |
| MPA | 586 |  | Q19 |  |  | Complement the text, insert a new section to clarify, as it is important to stress that importers and distributors primarily should inform the manufacturer. |  | Importers and distributors shall immediately inform the manufacturer and the authorized representative about suspected incidents and concerns related to the device. | Partially accept. Text revised and reference to guidance document on importers and distributors is included. |
|  | 601 |  | Q20 |  | Ed | Grammar |  | , an FSCA | Accept |
| FR | 149 |  |  |  | ed |  | Proposal to delete CJD as too limiting an example and add the proposed change | exposure to a high risk (e.g. cancer) of a significant number of the population, of a category of patients (diabetics, cardiac patients, etc.) or of a sensitive population (children, pregnant women, etc.) | Accept. Included as additional proposal. |
| FR | 540 |  | 16 |  | ed |  | add | The coordinating competent authority shall be the competent authority of the Member State in which the manufacturer has its registered place of business. However, if the MD isn’t put in his market, it could be another one. | Accept |
| FR | 2 |  |  |  | ed |  | Given that the corresponding articles of the regulation (UE) 2017/746 on IVDR are outside the scope of this document : add “for Medical Device” in the title | Questions and Answers on Vigilance Terms and Concepts for Medical Devices | Accept. |
| FR | 24 |  | 1 |  | ed |  | Add the exact definition of the incident in the MDR before explaining it; in fact, it is important to understand the elements that are added to the flowchart, knowing that the notion of undesirable side effects is new to the MDR. | An incident is defines in Article 2(64) of the MDR as “any malfunction or deterioration in the characteristics or performance of a device made available on the market, including use-error due to ergonomic features, as well as any inadequacy in the information supplied by the manufacturer and any undesirable side-effect;”. | Accept |
| FR | 31 |  | 1 |  | ed |  | Delete “i.e.”, given that the 3 points of article 2.65 are written. | A serious incident which is defined in Article 2 (65) of the MDR is an event where, a device malfunction or deterioration directly or indirectly resulted in or might result in/have resulted in the death, temporary or permanent serious deterioration in the state of health of a patient, user or other person or posed a serious public health threat. | Accept |
| FR | 74 (flowchart) |  |  |  | Te |  | The concept of incident includes undesirable side-effects; the same is true for serious incident. However, Article 87 requires a different reporting for serious adverse side effects. In fact, serious incidents are the subject of an MIR, except for expected side effects documented in the information and quantified in the technical documentation, which are the subject of a trend reporting. Only serious adverse side effects that are not documented and quantified in the technical documentation are subject to a MIR. Manufacturers shall report any statistically significant increase in the frequency or severity of non-serious events or expected adverse side effects in accordance with Article 88. | Left square : after the 2 paragraphs “If it is an adverse side effect not documented nor quantified in the technical documentation: it is the subject of an MIR. If it is an adverse side effect documented and quantified in the technical documentation: it is the subject of a trend reporting.” Right square : after the second paragraph “If it is an increase in frequency or severity of a documented and quantified serious adverse side effect or non-serious incident, therefore the manufacturer sends a trend reporting (art 88).” | Including this information may make the figure more complex. Side effects are addressed in question 7 and will be addressed in a separate guidance document currently under development by the TF for Trend reporting. Drawing a reference to ‘adverse side effects’ can also cause confusion as these are typically referenced in the context of clinical investigations. |
| FR | 162 |  |  |  | Ed |  | Add a reminder for manufacturers of the deadlines to which they must comply. | “Finally, manufacturers shall in all cases defer within the time limits specified in paragraphs 87.3 4 and 5” | More clarification needed of whether this refers to the last paragraph of c or when assessing if the criteria (A-C) are fulfilled. You are welcome to include the proposed text directly in the document, once it is send for WG endorsement. Question 11 and question 12 both outline the timeframes, which the manufacturer has to comply with. |
| FR | 296 |  | 7 |  | Ge |  | Add a reference to the flowchart. | See flowchart above. | As information on side-effects is not included in the flowchart (see also comment related to line 74), this reference is not included. |
| FR | 326 and 327 |  | 9 |  | Ed |  | Delete (1), (2) and (3). | “based on known information in the technical documentation, clinical or medical plausibility or investigative methods taken, cannot be excluded as a contributory cause of the serious incident.“ | Accept |
| FR | 648 (annex I) |  |  |  | Te |  | How is the reporting done with a trend report (MIR in addition?). | - | Annex deleted |
| AT | 64-68 |  |  |  | ed |  | Graphic not clear. Yes/No decision does not fit to the questions above (bullet points). We suggest to delete the bullet points in line 66-68. In our point of view it is much more clearer to start with the question whether the present complaint or report meets the definition of an incident according to Article 2 (64) MDR, followed by the question whether it fulfil the cirteria of a serious incident according to Article 2 (65) MDR | Please see alternative graphic from AT, which is attached to the email (name of document: “Reporting Criteria_AT” | See updated flowchart. |
| AT | 72 |  |  |  | ed |  | Typo in grafic: The manufacturer have to comply with the requirements for… | The manufacturer has to comply with the requirements for… | Accept |
| AT | 73 |  |  |  | ed |  | Typo: The manufacturer have to record the… | The manufacturer has to record the… |  |
| AT | 129 |  |  |  | ed |  | A serious deterioration in state of health of a person, as defined in Article 2(58) of the… | A serious deterioration in state of health of a person, as defined in Article 2(58) (b) of the | Task Force agreed to remove the reference to SAE definition to avoid any possible confusion |
| AT | 132 |  |  |  | ed |  | The wording “permanent” is missing | permanent impairment of a body structure or a body function | Accept. |
| AT | 138 |  |  |  | ed |  | Since it is an example and not limited to post-traumatic stress disorder, we should write e.g. instead of i.e. | …trauma (e.g. post-traumatic stress disorder (PTSD)). | Accept |
| AT | 132 |  |  |  | ed |  | The wording “permanent” is missing | permanent impairment of a body structure or a body function | Accept |
| AT | 129-145 |  |  |  | ed |  | Please refer to definition in MDR, since the change in wording leads to a change in the meaning of the definitions. | In our point of view the definition in the MDR is very clear. Old Q&A documents didn’t include extra definitions. Why do we don’t want to reference to the definitions in the MDR. To avoid typos, change of wording and misinterpretation we highly recommend to reference to the definitions in the MDR and NOT to include extra definitions in this guidance document! | Accept. Wording aligned with the definitions |
| AT | 147 |  |  |  | ed |  | Please include following term: (non-exhaustive list) | Examples of a serious public health threat can include (non-exhaustive list): | Accept |
| AT | 160-162 |  |  |  | ed |  | Just because an incident is complex, it does not necessarily mean that each of the medical devices involved must be related to the incident. We do not believe that it makes sense for medical device manufacturers to report complex incidents if they can safely rule out the possibility that their products were involved in the complex incident. | In complex situations, the manufacturers should err on the side of caution during assessment whether the device may have or potentially could have contributed to the serious incident. | Wording changed. |
| AT | 241-248 |  |  |  | ed |  | We do not understand the added value of this graphic. Maybe the description of the figure is satisfying. |  | Accept – it is deleted. |
| AT | 254 |  |  |  | ed |  | Please include chance in …and intentional but predictable misuse should… | …and abnormal use should… | Text deleted. |
| AT | 383-398 |  |  |  | ed |  | We suggest to change the year in this example from 2021 to 2022. |  | Accept |
| AT | 425-432 |  |  |  | ed |  | We suggest to change the year in this example from 2021 to 2022. |  | Accept |
| AT | 445 |  |  |  | ed |  | We suggest to write shall instead of should | explanatory statement shall be provided | Wording changed to ‘must’ |
| AT | 469 |  |  |  | ed |  |  | including device update or version rollback | See proposal |
| AT | 533-535 |  |  |  | ed |  | It could be helpful to include an example for better understanding | …FSN shall be consistent in all Member States (e.g. a translation error in the IFU that appears only in certain languages) | Accept |
| AT | 590-593 |  |  |  | ed |  | We suggest to divide the sentence into two sentences. | In principal, a device is considered to present a serious risk when, the combination of the probability of occurrence if a hazard causing harm and the severity of that harm is considered to require rapid interventional by the competent authorities. This should be based on the risk assessment, including the normal and foreseeable use of the device. | For simplicity, this text is removed and only proposal provided by MTE is included in the text. |
| AT | 617 |  |  |  | ed |  | Footmark Nr. 25: Doc “Questions and Answers on Articles 13 & 14 of Regulation (EU) 2017/745 and Regulation (EU) 2017/746” has been published. Link can be included. | https://ec.europa.eu/health/system/files/2021-12/mdcg_2021-27_en.pdf | Accept |
| AT | Annex I |  |  |  |  |  | General comment: Please change word “compliant” to “complaint” Under Criteria A: Simplify the sentence to: “An incident occurred” |  | Annex deleted |
| AT | 652-653 |  |  |  | ed |  | In our point of view arrows to Criteria B and C are not necessary. |  | Annex deleted |

| BfArM/DE | 211 | 5 | ge /te | This section does not describe the concept use error sufficiently as it misses use errors due to ergonomic features. Those use errors can also lead to lack of action, which shall not be considered as misuse. The definition in the ISO 14971 intentionally does not use the term “misuse” in conjunction with use error. We strongly suggest to delete the concept of a link between misuse and use error. According to the Oxford Dictionary misuse is” the act of using something in a dishonest way or for the wrong purpose”. User error can not be included in this purpose. What are “different factors occurring during normal use”? The user interface shall be considered as ergonomic feature. | Use errors can be described as an unintentional misuse of a device. These errors can be a result of ergonomic features, a mismatch between different factors occurring during normal use of a device such as; the characteristics or profile of the user, the user interface, or the nature of the task and/or the environment or constellation in which the device is used. | Text revised and changed. |
| --- | --- | --- | --- | --- | --- | --- |
| BfArM/DE | 220 | 5 | te | Although it is clear that the sentence, beginning with line 220, wants to distinguish between a use error and abnormal use, the chosen formulation leads to an unjustified global pre-determination of the causes of use errors, neglecting the possibility that a use error could be caused by ergonomic features. | - While use errors are typically a result of an In contrast to the unintentional misuse of a deviceuse error, ‘abnormal use’ is an intentional act or omission… | Text revised. |
| BfArM/DE | 232 | 5 | te | We think the line between normal and reckless use is too wide to use that as an explicit example. | Delete the example (line 232) | Accept |
| BfArM/DE | 254 | 5 | te | We are questioning the usefulness of this figure and we believe, that it does not describe the situation correctly. Reasonable foreseeable use is from our understanding not a concept used in the medical devices industry. It should not be mixed with the concept of reasonable foreseeable misuse. The first sentence in the description does not fit to figure 2. Figure 2 considers user error as a part of the normal use. The explanation makes a link between misuse and use error, which is not correct (see the above comments on use error) | Delete figure 2 and its explanation. | Text in question 5 revised and figure deleted. |
| BfArM/DE | 267 | 6 | te | Why does manufacturing belongs to ergonomics? | Delete “manufactured and” or replace by “contructed and” | Manufacturing deleted. See revised text |
| BfArM/DE | 267 | 6 | te | Important ergonomic features and sources of use errors are misleading alarms and software menue. They should be added as examples. | ‘Ergonomic features’ can be described as the design of the device and the way in which it is manufactured and constructed, including factors such as; physical features, and any measurement, monitoring or display scale, alarms or software menue. | Text changed. |
| BfArM/DE | 272 | 6 | ed | A mismatch or inconsistency between; the device features, … | Delete the semicolon after “between” |  |
| BfArM/DE | 272-273 | 6 | te | Information will not only be provided with accompanying information but by software responses and alarms, too. | A mismatch or inconsistency between; the device features, its intended purpose, performance, and the information provided e.g.in the accompanying information or alarms, | Accept |
| BfArM/DE | 275-276 | 6 | te |  |  |  |
| BfArM/DE | 280 | 6 | te | According to the MDR, use errors due to ergonomic features are reportable. What are user errors due to device design, that are due to ergonomic features but not device design related. | Delete “or device design related” or provide example(s). | Accept. Text revised. |
| BfArM/DE | 280 | 6 | te | With respect to the reporting conditions the wording of the MDR shall be used. | Any event related to a use error caused by ergonomic features or the device design, which lead or can potentially lead directly or indirectly led, might have led or might lead to the outcome | Accept. See revised text |
| BfArM/DE | 281 | 6 | te | Serious incidents must be reported to the CA (Art. 87) and must be investigated by the manufacturer (Art. 89). Therefore, the document should use “shall” instead of “should”. | Replace should with shall. | The word ‘must’ inserted to indicate this is a requirement. |
| BfArM/DE | 286 | 7 | te | The headline does not fully reflect the content. The last para of that section talks about reportability. That should be reflected in the headline. | 7. What is a ‘side-effect’ and how they are reported within the vigilance system? | Accept |
| BfArM/DE | 293 | 7 | te | We recommend reconsidering the philosophy in the document regarding "undesirable" and "expected" side effects. The MDR screwed it up as they dud not provide definitions if the different use was intentionally. We recommend using undesirable and unintended synonymously and linking the obligation to report to the "expected". | Delete lines 299 to 302 and replace them by: Expected side effect are subject to trend reporting according to Art 88. Unexpected side effects are in accordance with Article 87 reportable as single serious incident if they directly or indirectly led, might have led or might lead the outcome described in the definition of serious incident. | See revised text. |
| BfArM/DE | 334 | 9 | te | What is meant wit “known information in the technical documentation”? What is expected to receive? | Clarification necessary. | Accept. see proposals |
| BfArM/DE | 336 | 9 | ed | Correct wording? | Replace “investigative methods” with “investigation methods” | Accept |
| BfArM/DE | 349/350 | 10 | te | We cannot see a difference between the definition given in lines 342 to 347. | Replace lines 349/350 with The phrase ‘without undue delay’ should be interpreted the same way than ‘immediately’. | Accept |
| BfArM/DE | 421 | 12 |  | It shall not be an option for the manufacturer to explain the differences between the two dates, it shall be an obligation. | The manufacturer can shall explain the difference between the two dates | Accept |
| BfArM/DE | 442 | 13 | te | From BfArM’s experience, some manufacturers tend to move towards “non-reportable” on the basis of missing medical information. However, the conclusion “non-reportable” is to be derived from the investigation result of the device in question that either no incident in the meaning of Art 2 (64) or no outcome outlined in Art 2(65) has occurred or could occur. “Non-reportable’ should be the result of the manufacturer’s investigation of the device in question and not the pre-determination of the investigation. | When the manufacturer’s analysis according to Art 89(1) MDR of additional information, received after the submission of an initial MIR, conclude that the event is not a serious incident, then the box for ‘non-reportable’ in section 1.2 (e) of the MIR form shall be selected. We propose to add the following sentence: If the manufacturer is not finished with its investigations (including all started CAPAs), neither a final report nor a final, non reportable MIR should be send. | Accept |
| BfArM/DE | 449 | 13 | te | The form in which the manufacturer should provide his statement according to Art 87(11) MDR, third paragraph, is not described in the MDR. The MIR shall be used as well for such statements, marking the MIR with “final-non-reportable”. | We propose to add the following sentence: In principle, for statements according to article 87(11), third paragraph, the MIR shall be used as well for such statements, marking the MIR with “final-non-reportable”. | Accept. See revised text |
| BfArM/DE | 446-449 | 13 | ed | With reference to Art 87(11) third paragraph, the Q&A guidance document should use terms that are not weaker than those of the MDR. | Where the manufacturer of the device considers that the incident is not a serious incident or is an expected undesirable side-effect (which will be covered by trend reporting in accordance with Article 88 of the MDR) a rationale or an explanatory statement should shall be provided by filling out section 4.2(b) of the MIR report | Accept |
| BfArM/DE | 612 | 22 | te | Art. 88 MDR contains the sentence "any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side-effects that could have a significant impact on the benefit-risk analysis (...) that are unacceptable when weighed against the intended benefits." | A ‘significant impact on the benefit-risk analysis’ is whenever the increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation. The reportability of a trend results if the manufacturer’s update of the benefit-risk analysis reveals unacceptable risks when weighed against the intended benefits. An example of this can be… . | Accept. See revised text |
| BfArM/DE | 620 | 22 | te | Unacceptable risks suggests the consideration of a FSCA. A report according to Art 88 without a FSCA seems hardly possible. A text to this effect should be added from line 620. | A trend report according to Art 88 MDR shall include a disclosure of all planned and/or conducted associated corrective and preventive actions, otherwise, a rationale of why corrective actions cannot be carried out shall be provided. | The Task Force on trend reporting are addressing this point in the Q&A, which they are currently finalising. Therefore, members of the TF suggested that this is not included under Q 22. |
| DKMA | 24 | Q1 |  | The sentence “An incident is defined…” does not add value. | Delete “as an event….on the market”. Or as alternative add the full definition. | Accept |
| DKMA | 82 | Figure 1 |  | Yellow box: Clarification of date numbering is needed. | Add: Day of awareness = day 0. | Flowchart revised |
| DKMA | 162 | Q1 |  | The sentence “…the manufacturer should err on the side of caution” would benefit from a more clear wording. | “…the manufacturer should tend to issue a report on the side of caution” or “…the manufacturer should issue a report on the side of caution” | Accept. see revised text |
| DKMA | 252 | Figure 2 |  | “As errors can also occur during normal use of a device, both correct use and use errors are considered to be part of misuse during normal use of a device”. We are not sure how this should be understood. How can correct use be misuse? | Propose re-wording but we cannot provide an exact proposal as we are not sure what is meant. | Figure 2 deleted |
| DKMA | 274 |  |  | “call” must be a typo’. | Exchange with “can” or “may” | Accept |
| DKMA | 386 | Q11 |  | “..manufacturer determined that the complaint was not a serious incident” does not exclude death or public health threat. | Add “(no death or public health threat)” | Text revised |
| DKMA | 495 | Q14 |  | We are not sure of the use of the word “inkling”. Possible typo? | Exchange with “including”? | Accept |
| DKMA | 529-536 | Q16 |  | We have previously commented; to avoid several FSN drafts being circulated to different CA’s, we suggest that the evaluating competent authority is the CA of the MS in which the manufacturer or authorized representative is located. | We understand that it is not outlined in the MDR and will be addressed in the vigilance coordination procedure, which is under development. However, we propose that it is then not addressed in this document (The Q&A) as it is not clear what will be the conclusion in the vigilance coordination procedure. | The revised text is aligned with wording in MDR and does not contradict information provided in the vigilance coordination procedure. |
| DKMA | 613 | Q21 |  | “the root cause has been identified or a FSCA implemented”. Often a cause has been established before an FSCA is issued. | Propose exchange “or” with “and/or” | Accept |
| DKMA | 648-658 | Annex I, figure |  | Figure is not clear. We miss yes/no on the arrows and a box for the conclusion “do not report” | Add arrows and a box for if one or more of the A-C criteria is not met. A suggestion of the wording in the “No”-box is; “No obligations to report, but incident should be included in the post-market surveillance of the device and reevaluated if new information becomes available". Additional proposal for clarification; Merge blue and purple boxes for each criteria since the text in the purple boxes are examples and explanatory text to the blue boxes. They can thereby be read in connection. | Annex removed |
| EOF / GR | 379-380 | Q 11 | ge/te | We are proposing this change, because it is also possible that additional information (other than death / serious deterioration in the state of health) is received by the manufacturer, which is sufficient to make the event reportable. 1st alternative/proposed change would cover all cases and align with the Example that follows the specific paragraph. 2nd alternative/proposed change would specifically refer to the case that a reportable incident would need to be reported within a shorter timeframe, for simplicity. [The change of reportability (from non-reportabe to reportable) has already been explained in lines 351-356, anyway.] | “In case the manufacturer has first been made aware of an incident, and concluded it as a non-reportable incident was not reportable, and later is made aware of additional information, according to which he concludes that it is reportable, and subsequently is made aware of death or an unanticipated serious deterioration in a the person’s state of health, the timeframe with the earliest date of reporting must be applied.” OR “In case the manufacturer has first been made aware of a serious incident, and concluded it as a non-reportable incident, and later is made aware of death or an unanticipated serious deterioration in a person’s state of health, the timeframe with the earliest date of reporting must be applied.” instead of: “In case the manufacturer has first been made aware of an incident, and concluded it as a non-reportable incident, and later is made aware of death or an unanticipated serious deterioration in a person’s state of health, the timeframe with the earliest date of reporting must be applied.” | Accept. See revised text |
| EOF / GR | 404 | Q 12 | ge/te |  | “initial reportability assessment” instead of “initial device assessment” | Accept |
| EOF / GR | 439 | Q 13 | ed | Typo | “When the manufacturer’s analysis of additional information, received after the submission of an initial MIR, concludes that the event is not a serious incident, then the box for ‘non-reportable’ in section 1.2 (e) of the MIR form shall be selected.” | Text revised |
| EOF / GR |  | Footnote 25 | ge/te | Document published | Update to include the reference of the published document (MDCG 2021-27) | Accept |

| HPRA | 9-12 | Intro | Ed | The following sentence has a number of parts and could benefit from splitting out into two sentences for ease of reading. ‘To that end, this document is intended as a supporting tool for use by competent authorities, economic operators and other relevant parties to be read, in conjunction with the MDR, to establish a common understanding of the MDR vigilance requirements and obligations.’ | Suggested update: ‘This document is therefore intended to be a supporting tool for use by competent authorities, economic operators and other relevant parties. To establish a common understanding of the MDR vigilance requirements and obligations the document should be read in conjunction with the MDR’ | Text divided in two. |
| --- | --- | --- | --- | --- | --- | --- |
| HPRA | 28-30 | Q1 | Te/ed | Reference to ‘Incidents that do not fulfil the additional outcome requirements of serious incidents’ may cause confusion as it does not reference the potential to lead to one of these outcomes. | Suggested update to make this clearer.: ‘Incidents that have not led or could not lead to one of the outcomes described in section a, b, or c of Article 2(65) of the MDR are not reportable, but should be recorded and considered as part of the manufacturer's continuous post-market surveillance and should be reported in accordance with Article 88 of the MDR.’ Also made some minor editorial edits to included reference to ‘recorded and considered’ | Accept. Text revised |
| HPRA | 37 | Q1 | Te | Reference to ‘serious nature’ may cause confusion | Suggested update: ‘If an event has not led to the death, temporary or permanent serious deterioration in the state of health of a patient, user or other person or posed a serious public health threat, special consideration should be made to investigate…………..’ | Accept. |
| HPRA | 40-42 | Q1 | Te | Proposed update to make clear that a serious incident report needs to be submitter in case any confusion re reporting of an event | Suggested update: ‘If the manufacturer cannot exclude that the outcome might have led to a serious incident, the event should be reported to the relevant competent authority through submission of a serious incident report.’ | Accept, however adapted to avoid using the word ‘event’ and accommodate comment from MTE on consistency for using terms ‘incident, event, complain/report’ |
| HPRA | 45-46 | Q1 | ge | Would figure 1 be better under q2 as q2 introduces the three reporting criteria and this is covered in fig1? For part 2 of this figure, it states ‘what is the relationship between the incident and the manufacturer's device.’ | For consideration Suggested update to part 2: ‘what is the causality relationship between the incident and the manufacturer's device.’ | See revised flowchart, which aims to draw a clear parallel to question 2 without overlapping. |
| HPRA | 70 | Q1 | Te | Include reference to ‘must’ as it is a requirement. | Suggested update: Update to ‘Manufacturer must provide a Manufacturer Incident Report….’ | It is not an MDR requirement to use the MIR form. We therefore suggest wording in FN 9 |
| HPRA | 71 | Q1 | te | Include reference to ‘must’ as it is a requirement. | Suggested update: Update to ‘The manufacturer must comply with the requirements……’ | Accept |
| HPRA | 73 | Q1 | Te | Include reference to ‘must’ as it is a requirement. | Suggested update: The manufacturer must / have to record the report of | Accept |
| HPRA | 82 | Q1 | Ge | We found the yellow box here is a little bit confusing. We also thought it would be useful to reference Q11 and Q12 here. | Suggested update to amalgamate the two boxes in this section. Timeline in Calendar days*: Timeline in calendar days: - In principal: No later than 15 days. - In the event of a serious public health threat: No later than 2 days. - In the event of death or unanticipated serious deterioration in a person’s state of health: No later than 10 days. * Period begins on the day after the day of awareness of the potentially reportable incident. Further guidance on the application of these reporting timelines and day of awareness are provided in Q11 and Q12. | See revised flowchart |
| HPRA | 101-103 | Q2 | Te | There will not always be an instructions for use so proposed update to instead refer to intended purpose | Suggested update: ‘any malfunction or deterioration in the characteristics or performance of the device, i.e. a device that cannot/is losing its ability to function as intended and in accordance with the manufacturer’s intended purpose’ | Accept |
| HPRA | 129 | Q2 b | Te | As article 2(58) is included in SAE definition, would it be better to remove reference to the article to avoid any possible confusion? | For consideration | Accept |
| HPRA | 140-145 | Q2 b | ge | Note that this wording is almost verbatim to MDR definition. MDR definition includes reference to ‘imminent’ and ‘prompt’ while in the guidance these words are replaced with ‘immediate’ and ‘fast’ | Would it be simpler to use the exact MDR wording? | Accept |
| HPRA | 148-149 | Q2 b | te | Consider that a third more generic example would be useful here | Suggested update: Add: ‘Events that are of significant and unexpected nature such that they become alarming as a potential public health hazard, such as the possibility of multiple deaths occurring at short intervals’ This wording is amalgamated wording from Meddev 2.12/1 rev 8 | Accept |
| HPRA | 160-162 | Q2 c | ed | Consider splitting into two sentences for ease of reading | Suggested update: ‘In complex situations, it should be assumed that the device may have or potentially could have contributed to the serious incident. The manufacturer’s evaluation should err on the side of caution.’ | Accept. See revised text |
| HPRA | 172-174 | Q3 | ge | The meaning of this sentence is not clear to us | Consider the need to update to clarify | Text changed |
| HPRA | 208-211 | Q5 | ge | The meaning of the second sentence is not clear to us, in particular the reference to the ‘characteristics or profile of the user’ and ‘constellation in which the device is used’ | Consider the need to update to clarify | Text changed |
| HPRA | 213 | Q5 | te | Remove reference to ‘outcome’ in this sentence as it confusing when used in the same sentence as ‘might have led/might lead to’. Definition of a serious incident already covered in previous sections also | Suggested update: ‘Use errors which have led, might have led/might lead to a serious incident must be reported in accordance with Article 87(1) of the MDR’ | Text changed |
| HPRA | 229 | Q5 | te | We would remove reference to ‘Reckless use’ as what is/isn’t reckless use is quite subjective | Remove reference to ‘reckless use’ | Accept |
| HPRA | 242 | Q5 | te | We found Figure 2 is hard to follow. The text description is okay, however the illustration was a bit unclear | Consider updating to make clearer | Figure 2 removed |
| HPRA | 269-270 | Q6 | ed | Suggested updated text for ease of reading. | ‘Use errors caused by ergonomic features can lead to serious outcomes due to the unintentional nature of the error and the potential that the user is not aware of the occurrence of the error’ | Text revised |
| HPRA | 276-277 | Q6 | ed | Move ‘respectively’ to the end. | …..should be reported to the relevant competent authorities in accordance with Articles 88 or 87(1) of the MDR, respectively. | Text changed |
| HPRA | 295-296 | Q7 | te | Remove reference to ‘the health effect caused by the incident.’ Serious incident already defined earlier in the document so we suggest adding in ‘serious incident’ here rather than referring to outcome/potential outcome again | Suggested update: ‘………Article 88 of the MDR (not serious incidents) based on whether the incident meets the criteria of a serious incident or not’ | Text revised |
| HPRA | 309-311 | Q8 | te | ISO 14971 has moved away from reference to As Low as Reasonably Possible (ALARP) to As Far as Possible (AFAP) With regard to ‘is well known by the intended user’ – how will this be quantified? | Suggest updating to align with ISO 14971 Consider need to reference how ‘well known by the intended user’ will/can be quantified. | Accept |
| HPRA | 325-328 | Q9 | ed | This sentence was not clear to us. | Suggested update: ‘The manufacturer shall report in accordance with the requirements outlined in Article 87(3) of the MDR if the device cannot be excluded as a contributory cause of the serious incident. Determination of whether the device can or cannot be excluded as a contributory cause can be based off, but is not limited to the following: - known information in the technical documentation - clinical or medical plausibility - investigative methods taken ‘ | Text revised based on comments from MTE. |
| HPRA | 351-357 | Q11 | ed | Some suggested editorial changes for ease of reading | Suggested new text: ‘In the exceptional cases/situations where a manufacturer initially determines that an incident does not meet the criteria for reporting a serious incident, and then some time later the manufacturer obtains new information that impacts or changes the initial reportability assessment and and which results inthe reporting criteria of a serious incident being met, the period or reporting this case begins, at the time when the manufacturer has determined that the incident (e.g. event or complaint) fulfils the requirements for a serious incident (Article 2(65) of the MDR). In both situations i.e. general rule and the exceptional case/situation, the period ends on the 2nd, 10th or 15th day thereafter. The date that needs to be considered when reporting serious incidents,is the date when the manufacturer is made aware of the incident occurrence and not after it has conducted its investigation. | Accept. See revised text |
| HPRA | 368 | Q11 | ed | Swap ‘was’ with ‘is’ | ‘If the consequence of the serious incident is death or an unanticipated serious deterioration…….’ | Accept |
| HPRA | 379-381 | Q11 | Ed/te | Some suggested editorial changes. The current wording refers to death / unanticipated serious deterioration in state of health as the scenarios where a manufacturer would need to ensure the correct reporting timeline when the incident was first considered non reportable. However in such a scenario, the manufacturer would need to ensure correct reporting timeline for all reportable serious incidents rather than just death / unanticipated serious deterioration in state of health. | Suggested update: ‘In the case where the manufacturer is first made aware of an incident, and concludes that it as a non-reportable incident, and later is made aware of information that results in the manufacturer concluding that the incident is a serious incident, the appropriate timeframe which ensures the earliest date of reporting must be applied.’ | See revised text |
| HPRA | 383-387 | Q11 | ed | Editorial changes | Suggested update: ‘A manufacturer receives a complaint on 1st June 2021. The manufacturer determines that the criteria of a serious incident was not met and therefore does not submit a MIR to the relevant competent authority. The manufacturer subsequently receives additional information on 1st July 2021. Upon review of this information, the manufacturer determines that the complaint is in fact a serious incident: a MIR must be submitted at the latest by 16th of July 2021.’ | Accept |
| HPRA | 423-430 | Q12 | ed | Some suggested editorial updates. | ‘A manufacturer received a complaint on 1st July 2021. The manufacturer determines that the requirements for a serious incident are not fulfilled and does not submit a MIR to the relevant competent authority. The manufacturer subsequently receives additional information on 1st August 2021 and upon review of this information, the manufacturer determines that the complaint is a serious incident. The manufacturer should then submit a MIR within timelines outlined under the applicable sections of MDR, i.e., immediately and not later than 2, 10 and 15 days (respectively Article 87 (4), (5) and (3) of the MDR), from 1st August 2021. In the MIR form the manufacturer should insert the following dates in the fields: ‘ | Accept |
| HPRA | 495 | Q14 | ed | Presume ‘inkling’ was intended to be ‘including’? | For consideration | Accept |
| HPRA | 503-504 | Q14 | Ed/te | Change ‘to’ to ‘for’ and include reference to ‘communicated’ also. Transmitted is used in reference to Eudamed in MDR while definition of FSN refers to it as a communication. | Suggested updated text: ‘The FSN should be communicated/transmitted without delay for the attention of users of the device in question.’ | Accept |
| HPRA | 512-514 | Q15 | Ed/te | We think the wording here could be edited to make clearer. We also think the example of a software malfunction might not be the best example for a recall as usually these are fixed by a software upgrade. | Suggested updated text: ‘An example of such an FSCA would be where a recall of a device has taken place in a third country due to an identified malfunction with certain lots. If the lots affected by this recall have also been made available on the Union market, then all relevant competent authorities must be notified of the FSCA. ‘ | Accept |
| HPRA | 520-523 | Q16 | Ed/te | Reference to ‘corresponding FSCA’ may add confusion as evaluating CA will be for FSCAs even where a serious incident did not take place in the territory of a particular MS. | Proposed update: ‘The ‘evaluating competent authority’ is the national competent authority of the Member State responsible for the evaluation of the risks arising from reported serious incidents that occur within its territory and/ or of the adequacy of any FSCA(s) envisaged or undertaken by a manufacturer, in accordance with Article 89(3) of the MDR.’ | Accept |
| HPRA | 529-532 | Q16 | Te | We note that the current wording outlines that manufacturers should submit the draft FSN to all evaluating authorities and allow each evaluating CA to make comments on the draft FSN. Is it envisaged for art 89(8) that the draft FSN will be primarily reviewed and commented on by the coordinating CA? Or is it the expectation that the FSN will be reviewed by and commented on by all of the evaluating CAs? If the latter, this could slow issuance of FSNs down significantly as for a FSCA that affects 20 MS’s then the manufacturer could be receive comments from 20 evaluating authorities. | Consider updating if the current wording does not align with the intention. | Accept. See revised text. |
| HPRA | 537-540 | Q16 | Te/ed | Will be linked to the answer to the above, however we have suggested an update here. | Suggested update: ‘In a case where it is determined that the coordinating procedure of Article 89 (9) of the MDR should be initiated, the evaluating competent authority should inform the manufacturer that it has taken the role of the evaluating/coordinating competent authority in accordance with Article 89 (9) of the MDR either directly or through Eudamed.’ | Partially accept. See revised text. |
| HPRA | 586-593 | Q19 | ge | With regard to reference to rapid intervention by market surveillance authorities, we consider that it may be useful to clarify that rapid intervention by the MS authority may be to alert the manufacturer or to open an investigation. As it currently reads, it infers that it will be up to the MS authority to act upon notification of the issue rather than the manufacturer to act which goes against the MS articles. The second paragraph outlines how a serious risk can be identified and includes reference to basing this off the risk assessment / probability of occurrence etc. Importers/distributors will not have visibility to this information so this current wording may not be practical for importers/distributors. An importer is not going to carry out probability and occurrence of harm for estimating serious risk. Are the definitions of serious risk in the two paragraphs slightly inconsistent? The first paragraph defines it as a situation where a serious harm is likely to happen and requires rapid intervention by the MS, whereas the second paragraph defines it as where based on the risk assessment it is determined that a rapid intervention is required by the MS. | General comments for consideration. | First section addressed by adding examples of rapid intervention. With regard to risk assessment, it is clarified that this is performed by the manufacturer. With regard to importers and distributors, a sentence on their contribution is added. |
| HPRA | 621-623 | Q22 | ed | The first sentence in this question was a bit unclear to us, Suggested update provided. | ‘A ‘significant impact on the benefit-risk analysis’ is where an increase in the frequency or severity of incidents occurs, increases that are not categorized as serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation.’ |  |
| IT | 287 - 296 |  | GE/TE | Considering the ongoing discussion in Trend Report Task Force we propose to revise the text for Question 7. We believe that a side effect can be “expected” or “unexpected” and also “undesirable”. So, it could be possible to have an “expected-undesirable” side effect but also an “unexpected-undesirable”. The terms expected/unexpected are referred to the fact that a side effect is or not previously known and considered by the manufacturer in its risk/benefit evaluation ratio. The term “undesirable” is instead referred to the clinical characteristics of the side effect. From our point of view, an unexpected side effect represents a “new” situation for the manufacturer. For this reason, we think that all the unexpected side effects shall be reported as a serious incident (Mir Form) and after monitoring it for a while the manufacturer could define a risk/benefit, a threshold, and include it in the expected side effects and then it could be reportable in accordance article 88. Regarding this point please consider that in the art.88 for trend reporting the “expected undesirable” side effect should be considered. | A ‘side-effect’ is any unintended medical manifestation in the human body, including patient discomfort, as a consequence of the normal or abnormal use of a device. Side-effects are, in principal, not the result of a malfunction, deterioration in the device’s characteristics or performance, or an inadequacy in the information provided by the manufacturer. In the MDR are mentioned the “undesirable”, the “expected” and the “expected undesirable” side-effects. An undesirable side-effect, for the purpose of this guidance, could be expected or unexpected. Expected side-effects should be quantified in the manufacturer’s technical documentation and be clinically acceptable in terms of individual patient benefit, and clearly documented in the product information. Expected side-effects, should be reported in accordance with the requirements for trend reporting pursuant to Article 88 of the MDR. Unexpected side-effects which are, inter alia, events that were not considered or addressed by the manufacturer in its risk analysis, should be reported in accordance with Article 87 of the MDR (serious incidents) or Article 88 of the MDR (not serious incidents) based on the health effect caused by the incident. Only after the manufacturer has monitored it for a while, it will be possible to define a risk/benefit ratio and a threshold, and therefore include that side effect among the expected side effect and also it could be reportable in accordance to Article 88. | Accept. See revised text |

| PT | 293 | 7 | te | In our opinion undesirable is not directly related with events not addressed in the risk analysis by the manufacturer. The events not addressed by the manufacturer on risk analysis are those one that are unexpected. Undesirable is an event that was not intended by the manufacturer. It is possible to have undesirable side-effects expected (that are addresses in risk analysis) and undesirable side-effects unexpected (and these were not addressed by the manufacturer on risk analysis). “Undesirable side-effects which are, inter alia, events that were not considered or addressed by the manufacturer in its risks analysis should be reported in accordance with Article 87 of the MDR (serious incidents) or Article 88 of the MDR (not serious incidents) based on the health effect caused by the incident.” | Undesirable Unexpected side-effects which are, inter alia, events that were not considered or addressed by the manufacturer in its risks analysis once they were not foreseen/expected. should be reported in accordance with Article 87 of the MDR (serious incidents) or Article 88 of the MDR (not serious incidents) based on the health effect caused by the incident.” Undesirable side-effects are unintended effects/events by the manufacturer considering the medical purpose of the device. The undesirable side effects can include expected or unexpected side-effects. Undesirable side-effects should be reported in accordance with Article 87 of the MDR (serious incidents) or Article 88 of the MDR (not serious incidents) based on the health effect caused by the incident. | Se revised text. |
| --- | --- | --- | --- | --- | --- | --- |
| PT | 586-589 | 19 | te | “…a ‘serious risk’ is defined as a situation where a serious harm affecting patients, users or the public, as a result of use of a device, is likely to happen and require rapid intervention by the market surveillance authorities. A serious risk may include situations where the effects of the risk are not immediate.” We have some doubts if these situations of serious risk can include serious incidents. Vigilance is dealing with serious incidents; for this reason it seems to us that these are other situations besides serious incidents. MDR text refers “device presents a serious risk or is a falsified device”. | We will suggest the following text for clarification: “…a ‘serious risk’ is defined as a situation where a serious harm affecting patients, users or the public, as a result of use of a device, is likely to happen and require rapid intervention by the market surveillance authorities. A serious risk may include situations where the effects of the risk are not immediate. These situations defined as “serious risk” are other situations besides serious incidents. | The text has been revised and simplified to avoid confusion. |
| FI |  | Q. 21 what is a common and well documented serious incident |  | There are 3 different justification for PSR 1) Fsca released 2) root cause have been identified 3) common and well documented, so common and well documented can not be reasoned by an other criteria. Common means that there are number of incidents, but what is well documented. Is it enough that the incident type is known? |  | See revised text. Question reworded. |

Following elaboration provided for alignment with text in Article 2(64) of the MDR: a device malfunction or deterioration in the characteristics or performance of a device made available in the market, including use error due to ergonomic features, as well as any inadequacy in the information supplied by the manufacturer and any undesirable side-effect.

Expected side-effects are clinically well known, i.e. in the medical, scientific or technological field, or have been clearly identified during clinical investigation or clinical practice by the manufacturer. Expected side-effects should be clinically acceptable when weighed against the evaluated individual benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use.

The product information can include the accompanying documentation, such as the instructions for use and the labelling of the device, training material and the technical documentation.