Question and Answer on Trend Report

Introduction

This document aims to explain and clarify questions related to Trend reporting as outlined in Article 88 of the Regulation (EU) 2017/745 on medical devices (MDR) and Article 83 of Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR). A common understanding and clarification related to Trend reporting is deemed essential for an effective and harmonised implementation of the vigilance requirements under the MDR and IVDR.  To that end, this document is intended as a supporting tool for use by competent authorities, economic operators and other relevant parties to be read, in conjunction with the MDR and IVDR, to establish a common understanding of the MDR and IVDR vigilance requirements and obligations.

The requirements on regarding Trend Reporting, as defined in the Article 88 of the MDR, are applicable from the 26th May 2021.

The requirements on regarding Trend reporting, as defined in the Article 83 of IVDR, are applicable from 26th May 2022.

The Trend Reporting under the MDR requires a consistent, standardized and systematic review of all incidents and expected side effects apart from serious incidents, by medical device manufacturers. As far as IVDR requirements are concerned, a standardized and systematic review of all incidents, apart from serious incidents, and expected erroneous results is needed.

The requirements regarding trend reporting differs between the directive 93/42/EEC and MDR. However, trend reporting based on the MDR could be applied also to the old devices. Consequently it is not necessary to distinguish between ‘old’, ‘legacy’ and ‘MDR’ devices for the purpose of defining the vigilance reporting obligations of the medical devices.

The requirements on trend reporting are applicable to:

- Legacy devices (MDR and IVDR)

- MDR certified devices

- IVDR certified devices

- Old devices (MDR)

The Terms and Concepts, useful for Trend Reporting, as outlined in Section 2 of Chapter VII of the Regulation (EU) 2017/745 on medical devices, are defined on the document Questions and Answers on Vigilance Terms and Concepts (add reference when available).

The concepts defined in this document are applicable to the corresponding articles of the Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR), when an IVDR-specific Q&A document is not available.

- What shall a manufacturer report with a Trend Report?

The Medical Device Regulation 2017/745 Article 88 (1) requires manufacturers to report as Trend Report, to the Competent Authorities where the incidents occurred (art.92(6) ), any statistically significant increase in the frequency or severity of incidents, that are not serious incidents, or that are expected undesirable side-effects that could have a significant impact on the benefit-risk analysis.   The statistically significant increase shall be established in comparison to the foreseeable frequency or severity of such incidents in respect of the device, or category or group of devices, in question during a specific period as specified in the technical documentation and product information.

The In Vitro Medical Device Regulation 2017/746 Article 83 (1) requires manufacturers to report as Trend Report, to the Evaluating Competent Authority, any statistically significant increase in the frequency or severity of incidents or expected erroneous result established in comparison to the stated performance of the device as referred to in points (a) and (b) of Section 9.1 of Annex I and specified in the technical documentation and product information.

The Competent Authority shall be notified  in cases where a statistically significant increase  of incidents or expected undesirable side effects (MDR) or expected erroneous result (IVDR) is detected and that change could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. Trends should be identified by the manufacturer as they can be indicative for a change in the risk-benefit ratio.

- When a manufacturer should report incidents, expected undesirable side effects or expected erroneous results to a Competent Authority using a Trend Report?

The manufacturer has to record incidents and expected undesirable side effects in accordance with Articles 83-86 of the MDR prescription for the post market surveillance.

As specified in the Article 83 (2) of the MDR, a manufacturer in its post market surveillance system has to gather, record, analyze, actively and systematically, any relevant data on the quality, performance and safety of a device throughout its entire lifetime, in order to be able to draw the necessary conclusions and to determining, implementing and monitoring any preventive and corrective actions. As also specified in the Article 83 (3) of the MDR, data gathered by the manufacturer's post-market surveillance system shall, among other things, be used to detect and report trends in accordance with Article 88 of the MDR.

Whenever a statistically significant increase in the frequency or severity of non-serious incidents or of expected undesirable side effects or expected erroneous result is detected and results in a change of the risk evaluation and has a significant impact on the benefit-risk analysis, the manufacturer shall draw up a trend report as defined in the Article 88/83 of the MDR/IVDR.

It is worth mentioning that while assessing the increase in the events frequency and severity, it is implied that the risk management system must be based on quantitative categorization of probability of occurrence of harm and severity of harm versus the option to use either qualitative or quantitative categorization as required by ISO 14971: Medical devices - Application of risk management to medical devices.

A trend report can be drawn up when the manufacturer has performed a monitoring of incidents over time and has compared the severity and the frequency of the occurrences with the threshold defined during the design process and documented in the risk analysis.

- What happens if a NEW unexpected undesirable side effect or incident is identified?

Unexpected undesirable side-effects or newly identified incidents are events that were not considered or addressed by the manufacturer in its risk analysis should be assessed and

- reported in accordance with Article 87 of the MDR (serious incidents) and

- documented and analyzed during the update of the risk analysis and the risk management file.

The update of the risk-benefit assessment should take the identified new risks in consideration and labeling updates implemented as required.

Whenever a statistically significant increase in the frequency or severity of incidents or of unexpected undesirable side effects is identified and confirmed according the manufacturer internal procedures, the manufacturer should conduct a full assessment on  impact of the trend on the benefit-risk ratio of the product and shall draw up a trend report as defined in the Article 88/83 of the MDR/IVDR.

- How to link the PMS plan with the trend reporting process?

Taking into account the requirement of Post Market Surveillance (PMS) (MDR, Chapter VII: Post-market surveillance, vigilance, and market surveillance) manufacturers do not only need to perform trend reporting per Articles 88 and 83, but they must also include in the PMS Plan the methods and protocols to manage the events subject to trend reporting. This includes the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period (per Annex III, section 1.1 b: Technical Documentation on Post-Market Surveillance).

As per Article 88 (1) (MDR), the manufacturer shall specify, in the post-market surveillance plan referred to in Article 84 for MDR:

-  how to manage the incidents subject to the trend report;

-  the methodology used for determining any statistically significant increase in the frequency or severity of such incidents;

- the observation period.

Article 84 (MDR) specifies that the requirements for the post-market surveillance plan are set out in Section 1 of Annex III which says that “methods and protocols to manage the incidents subject to the trend report as provided for in Article 88, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period.”

As per Article 83 (1) (IVDR) the manufacturer shall specify, in the post-market surveillance plan referred to in Article 79 of the IVDR:

- how to manage the incidents subject to the trend report;

- the methodology used for determining any statistically significant increase in the frequency or severity of such events or change in performance;

- the observation period.

- How to link the trend reporting process to general Quality Management System?

The general requirements on the quality management system (QMS) of medical devices and IVD’s are described in section 4.1 of the ISO 13485 as well as in Annex IX and the article 10 of the MDR and article 10 of the IVDR. The Implementation of the quality management system shall ensure compliance with this Regulation and the applicable standards.

The trending process is one of the key elements of vigilance, risk management and PMS processes and of the QMS. The manufacturers of medical devices and IVD’s shall plan and implement the monitoring, measurement, analysis needed to demonstrate conformity, the performance, and the safety of their products and to ensure conformity and maintain the effectiveness of the quality management system.

Therefore, and as required by article 10.13 of the MDR and Article 10.12 of the IVDR manufacturers shall have a system for recording and reporting of incidents and field safety corrective actions as described in Articles 87 and 88 of the MDR and Articles 82 and 83 of the IVDR.

The risk management and clinical evaluation, PMS processes under which manufacturers operate are essential key elements in trend reporting being the basis of determining:

- when trends become significant,

- when the acceptable (lowest) risk level (severity-harm combination) might become threatened,

- when the risk-benefit analysis of residual risk has led or may lead to unacceptable risks to health or safety of patients,

- when the residual risk does not constitute an acceptable risk when weighed against the benefits to patient,

- when actions are necessary by manufacturers to address unacceptable risks.

- Which incidents/undesirable side-effects/expected erroneous result from a geographic perspective have to be included in Trend Reporting?

The manufacturer has to record incidents and expected undesirable side effects in accordance with Articles 83-86 of the MDR prescription for the post market surveillance.

In order to have a statistically significant data as a base for a trend report, the manufacturer has to take into account all the incidents related to a specific trend and from a geographic perspective has to consider all the incidents that have occurred worldwide. The worldwide data about incidents collected by the manufacturer during PMS activity, are all useful to define the statistically significant dataset which establish a trend. This trend be notified to the countries in which MDR is applicable and where related incidents occurred.

Therefore, the detection of trends is based on the data collected about incidents occurring in Europe and worldwide. Consider limitation to those countries where incident reporting is applied. The manufacturer should prepare a trend report when the trend includes at least one incident that occurred in the European Economic Area (EEA)+ Turkey (TR)+ Northern Ireland (XI).

- Does a trend report refer only to one device or potentially also to a group of devices?

As indicated in the Article 88 (1) of the MDR, a trend report could be performed for a single device as well as a category or group of devices if they are impacted by the same incidents/ undesirable side-effects/or expected erroneous.

For this reason, in the Section 2 of the Trend Report Form the manufacturer can select the device scope type choosing between the medical device high level EMDN code i.e. group of medical devices, Basic UDI-DI/Eudamed DI, UDI-DI/Eudamed ID and UDI-PI/Eudamed ID (lot/batch number). The old devices could be reported only by the level of device name.

Consider that only one choice is possible in the device scope type.

- Could the device scope and the scope of the trend be changed during the trend report process and the trend report life cycle?

When a trend report is performed and a device scope is defined it has to be kept the same during the trend report process and trend report life cycle. This is because a trend report is based on statistical data collected by the manufacturer with a specific methodology and they have to be referred to a specific device scope in order to avoid any confusion and to allow to the Evaluating Competent Authority to carry out the appropriate assessments on the data provided.

If a new device scope is identified for the same device/category/group of devices, it is necessary to create a new trend report with its own process and life cycle. If during the life cycle of the trend report, after further investigation of the manufacturer the device scope changes, it could be included in the trend report follow up ONLY if the assessment, under which the trend report is based on remain unchanged (for example a new batch of device is identified).

- If a device has different problems, how should they be reported?

If a device has different problems that meet Article 88/83 requirements, each problem should be reported in a single trend report.

- With reference to Article 88/83 of the MDR/IVDR and the criteria for trend reporting, what is meant by a ‘significant impact on the benefit-risk analysis’?

A ‘significant impact on the benefit-risk analysis’ is whenever the increase in the frequency or severity of incidents, that are not categorized as serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation (EN ISO 14971). A significant increase of the severity could be measured as possible consequence of the incident and / or increase of the probability of occurrence of harm in such a way that could lead to unacceptable Residual Risk and significant impact on the benefit-risk analysis. An example of this can be, a risk zone changes from low to medium, or and increase that is from medium to high, within the risk matrix as suggested in accordance with ISO 14971.

In line with Article 88/83 of MDR/IVDR and in reference to sections 1 and 5 of Annex I / Chapter I, the final decision to report is complemented by health hazard assessment in which the residual risk is weighed against the intended benefits.

- What is the “threshold” mentioned in the section 3 (b) of the Trend Report Form?

The established thresholds value for reporting refers to the threshold defined by the manufacturer in its risk assessment in a way to avoid over or under-reporting. In this section the manufacturer should describe how he has determined potential risk and which signal was detected.

Thresholds of risk for trend reporting should be defined at the beginning of the designing and manufacturing process in accordance with the risk-benefit analysis and reference to them should be reported in trend report.

In particular,

Risk can be considered as the results of two components:

  - probability of damage occurring;

  - consequences of the damage, meaning gravity of the damage.

Benefit is defined as the benefit of a medical device which is related to the likelihood and extent of the health improvement expected from its use.

Risk can be considered as the results of three components:

  - the expected performance during its clinical use;

  - the clinical outcome expected from this service;

  - factors relevant to the risks and benefits of other treatment options.

As the standard ISO 14791 do not define exactly the “acceptable risk” the evaluation is required by the manufacturer who should have a suitable system in place for proactive scrutiny.

Some qualitative methods can be identified in agreement with the EN ISO 14971.

Examples of matrix for qualitative and semi-qualitative risk assessment can be identified consulting the EN ISO 14971.

- Which Trend Report Form has to be used before and after Eudamed becomes fully functional?

Trend reporting via the electronic system (EUDAMED) referred to in Article 92 of the MDR and Article 87 of the IVDR using the Trend Report Form will be applicable only when Eudamed becomes fully functional.

As per art. 92 (6) MDR and art. 87 (6) IVDR the trend reports shall be automatically transmitted upon receipt via the electronic system to the competent authorities of the Member States in which the incidents occurred. It is advisable to submit the trend report also to the National Competent Authority (NCA) where the manufacturer or its authorized representative has his registered place of business.

In the meantime, the Eudamed is fully functional the MDR/IVDR Trend report requirements are applicable and the alternative administrative technical solutions have to be adopted as per guidance MDCG 2021-1. The current Trend report form should be used and it should be sent to the national vigilance systems of member states.

- Which are the documents related to the Trend Report when the Eudamed is fully functional?

When the manufacturer issues a Trend Report, they have to take into account 2 documents:

- Eudamed Trend Report Form. This document submitted directly in Eudamed including the following information: Member States where the incidents occurred, the administrative information, description of the devices and the detected trend (background, established threshold, methodology, observation period, root cause analysis, actions to be taken or already implemented to the reduce the risks for the users/patients).

- Trend Report document. Additional document should be added as required. This document is a pdf file attached to the Eudamed Trend Report Form and includes more detailed information about the affected devices, the method and description of the detected trend and its investigations and conclusions.

As mentioned above, the Eudamed Trend Report Form (and its related helptext) will be applicable only when Eudamed becomes fully functional.

- Can a Trend Report lead to the definition of a corrective action/preventive action?

During the assessment of the data of the trend report, the manufacturer can issue a corrective action including field safety corrective action/preventive action to prevent or reduce risk for patient or user of the device. In this context the manufacturer should notify a Final trend report and initiate an appropriate action (e.g. FSCA, CAPA, update of labeling, training).

If the evaluating CA is not satisfied with the action taken by the manufacturer, they could require additional corrective actions as appropriate as specified in Article 88 (2)/ 83 (2) of MDR/IVDR.

- Which kind of request Competent Authority can make after the assessment of a trend report?

The competent authorities, as specified in Article 88 (2), may conduct their own assessments on the trend reports and require the manufacturer to adopt appropriate measures in accordance with the MDR/IVDR in order to ensure the protection of public health and patient safety.

When a trend report is received the CA assesses it. This assessment in some cases could lead to a request of re-evaluation by the manufacturer.

the appropriate measures that a CA could request to a manufacturer can be of various types as long as all those included in the Regulation itself.  For example, it can be asked to issue a corrective/preventive action or any other kind of measure in order to ensure the protection of public health and patient safety.

- How can the manufacturer manage the Custom-Made Device CMD in the scope of Trend Report?

The same process related to the identification of a Trend applies also on the CMDs.

With the absence of stated exceptions, CMD manufacturers must meet nearly all of the MDR requirements.

As defined also in the guidance MDCG 2021-3, for risk management, post-market surveillance and clinical evaluation life cycle processes as defined by the MDR, CMD manufacturers should apply these obligations to groups of devices with the same intended purpose, materials used, process utilized, same principal design etc. and not to each individual CMD.

As a consequence, the requirements of Article 88 of the MDR apply to CMD’s.

- What are some examples for reporting a Trend Report?

In order to further clarify what shall be reported in a trend report, some examples are provided below:

Examples for MD:

- A trend of expected undesirable side effects causes a risk zone changes from low to medium, or an increase from medium to high, within the risk matrix as suggested in accordance with ISO 14971, Medical devices, Application of risk management to medical devices;

- A number of similar incidents, according to Article 2 (64) MDR, that are not serious incidents reported by different healthcare facilities shows a statistical increase with a daily or monthly frequency;

- An increase in an incident scenario that does not meet reporting requirements of Article 87 (harm severity: minor) causes a risk management threshold breach and leads to an increase in probability from “remote” to “occasional”

Examples for IVD:

- Increase of false positive or false negative results

- misdiagnosis

- delayed diagnosis

- delayed treatment

- inappropriate treatment

- transfusion of inappropriate materials

A distinction between ‘legacy’ and ‘old’ devices may remain relevant for the registration of the devices in Eudamed, i.e. how to link a serious incident report registered in the vigilance module to the registration of a device in the UDI / device module.  This issue has been addressed by MDCG guidance 2019-5.

referred to in Sections 1 and 5 of Annex I MDR and IVDR

EN ISO 14971 Medical devices, Application of risk management to medical devices.

Probability of occurrence of harm = the probability of hazard or hazardous situation multiplied by the probability of hazard or hazardous situation causing harm (EN ISO 14971 definition)

Residual Risk = combination of the probability of occurrence of harm and the severity of that harm (EN ISO 14971 definition).

MDCG 2021-1 “Guidance on harmonized administrative practices and alternative technical solutions until EUDAMED is fully functional

Refer also to MDCG 2021 – 3 guidance “Questions and Answers on Custom-Made Devices” Question 8 and 9.

Question 8 MDCG 2021 -3: 8. What are the obligations of CMD manufacturers according to the MDR?

With the absence of stated exceptions, CMD manufacturers must meet nearly all of the MDR requirements. Although the authorised person issuing the written prescription primarily determines the design and intended purpose of the CMD, it is the responsibility of the CMD manufacturer to consider which MDR Annex I requirements are applicable to the device at hand.

Additionally, and as any other medical device manufacturer, CMD manufacturers shall establish, document, implement, maintain, keep up to date and continually improve a quality management system (QMS) that shall ensure compliance with the MDR in the most effective manner and in a manner that is proportionate to the risk class and the type of device. The QMS must address all elements described in Article 10(9) of the MDR.

To implement an MDR compliant post-market surveillance system, the CMD manufacturer should establish appropriate communication channels with relevant healthcare providers/professionals or patients to receive feedback on the quality, performance and in particular the clinical performance and safety of the devices in the field.

For risk management, post-market surveillance and clinical evaluation life cycle processes as defined by the MDR, CMD manufacturers should apply these obligations to groups of devices with the same intended purpose, materials used, process utilised, same principal design etc. and not to each individual CMD.

In accordance with Article 87(1) of the MDR, CMD manufacturers shall report to the competent authorities any serious incidents and/or field safety corrective actions as soon as they learn of them.