## Taul1
| Section | Help text, concentrate on the trend report specific issues, those fields which are identical with the MIR will be according to the updated MIR and the structure of the form will be adjusted in the EUDAMED web interface. So far an open question is whether the current trend report form will be updated to enable reporting with the MDR, IVR and legacy devices. | Combined initial & final | Initial | Follow Up | Final
1. Administrative information
1.1 | Corresponding competent authority(ies)
a | Member State(s) in which the incidents expected undesirable side effects, expected erroneous results (IVD) occurred | Manufacturer thick/tick the EEA country(ies) where the reported events occured. Possible to select all countries with one click. he has received the incident reports and add manually list the countries outside EEA+TR+IX where same type of incident has oocured. The countries outside EEA+TR+IX should be added manually. | Y | Y | Y | Y
b | To which NCA(s) is this report being sent? | Ref number of NCAs to whiche the report is sentd auto populated by Eudamed | Y | Y | Y | Y
c | NCA where the manufacturer or authorised representative is located | Name and number of the NCA | N | N | N | N | Ask whether auto populated
d | Reference number assigned by EUDAMED for this incident/Expected undesirable Side Effects/ Expected erroneous results (IVD) EUDAMED Reference number for this MTR | The reference number assigned by Eudamed for this Trend report (after uploading/entering the trend report into Eudamed) | N | N | N | N
1.2 | Date, type, and classification of trend report
a | Date of submission | Date when you submit the report, autopopulated by Eudamed | Y | Y | Y | Y
b | Date the trend was identified | The date during the trending period, when the manufacturer detected the trend | Y | Y | Y | Y
c | Time period of trend analysis | The time trending period during which the trend was identified | Y | Y | Y | Y
d | Regulatory scope type MD (MDR, MDD, AIMDD) or IVD (IVDR, IVDD) | Select the device type, medical device or IVD-device, the MTR covers | Y | Y | Y | Y
d e | Basis for the report (type) and Basis for the report – what increased | Classify the trend report based on the cause for reporting: incident, expected side-effect (MD) or expected erroneous results (IVD) and whther it is based on the increased frequency or severity or to both. | Y | Y | Y | Y
f | Type of Report | Initial Report - Choose this for your first report on a particular trend. The trend report document is not mandatory with the initail report. Follow-up report - Choose this to provide additional/interim information during your investigation. You can submit more than one follow-up report for each trend. Note this option is only available once you have submitted an initial report. The Trend report document is mandatory. Combined initial & final - Choose this if you have available all the required information during for the details of the initial and final report within the timeframe for reporting time. Note that this option is not available once you have submitted an initial report. The Trend report document is mandatory. Final - This is your formal statement of the outcome of your investigation, including any corrective and preventive actions proposed or taken. The Trend report document is mandatory. | Y | Y | Y | Y
g | In case of initial and follow-up reports, please indicate the expected date of next report | The date by which you expect to be able to submit your next report on this trend. This may be either a Follow-Up or Final report. |  | N | N | N
h | The trend was detected within a PMCF/PMPF investigation? Yes/No | Indicate Yes/No if the trend was detected in PMCF/PMPF investigation
i | If the trend was detected within a PMCF/PMPF investigation; please provide the Eudamed ID of that PMCF/PMPF investigation | Please provide the Eudamed ID of that PMCF/PMPF investigation. The field is displayed in cases the previous answer was yes. |  | N | N | N
1.3 | Submitter information
1.3.1 | Submitter of the report information
a | Submitter of report | Identifies the actor who is submitting the report | Y | Y | Y | Y
b | Manufacturer's reference number for this trend report | The reference number assigned by the manufacturer | Y | Y | Y | Y
1.3.2 | Manufacturer information
a | Manufacturer Organisation name | The name of the Manufacturer for the device involved in this Trend report | Y | Y | Y | Y
b | Single Registration Number (SRN)/Actor ID | A SRN is the Single Registration Number that is issued through EUDAMED to manufacturers, authorised representatives and importers by the competent authority in accordance with Article 31 MDR and 28 IVDR. An Actor ID is the identifier issued to all actors registered in EUDAMED for their identification in the system. It will be automatically generated by EUDAMED and issued once the registration request is approved by the relevant competent authority. | Y | Y | Y | Y
c | Contact's first name | First name of the manufacturer contact person | Y | Y | Y | Y
d | Contact's last name | Last name of the manufacturer contact person | Y | Y | Y | Y
e | E-mail | E-mail of the manufacturer contact person . | Y | Y | Y | Y
f | Phone | Telephone number of the manufacturer contact person | Y | Y | Y | Y
g | Country | The country where the manufacturer is located. | Y | Y | Y | Y
h | Street | The street of the manufacturer | Y | Y | Y | Y
i | Street number | The street number of the manufacturer | N | N | N | N | N
j | Address complement | The address where the Manufacturer is located (e.g. building name) | N | N | N | N | N
k | PO Box | The PO box of the manufacturer | N | N | N | N | N
l | City name | The name of the city where the manufacturer is located. | Y | Y | Y | Y
m | PostcodePostal code | The postal or zip code where the manufacturer is located | Y | Y | Y | Y
1.3.3 | Authorised representative information
a | Authorised representative organisation name | The name of the Authorised representative for the device involved in this adverse incident trend report | Y | Y | Y | Y
b | Actor ID/SRN (Single Registration Number, SRN) | A SRN is the Single Registration Number that is issued through EUDAMED to manufacturers, authorised representatives and importers by the competent authority in accordance with Article 31 MDR and 28 IVDR. An Actor ID is the identifier issued to all actors registered in EUDAMED for their identification in the system. It will be automatically generated by EUDAMED and issued once the registration request is approved by the relevant competent authority. | Y | Y | Y | Y
c | Contact's first name | First name of the Authorised representative contact person | Y | Y | Y | Y
d | Contact's last name | Last name of the Authorised representative contact person | Y | Y | Y | Y
e | E-mail | E-mail of the Authorised representative contact person . | Y | Y | Y | Y
f | Phone | Telephone number of the Authorised representative contact person | Y | Y | Y | Y
g | Country | The country where the Authorised representative is located. | Y | Y | Y | Y
h | Street | The street of the Authorised representative | Y | Y | Y | Y
i | Street number | The street number of the manufacturer Autorised representative | N | N | N | N
j | Address complement | The address where the Authorised representative is located (e.g. building name) | N | N | N | N
k | PO Box | The PO box of the Authorised representative | N | N | N | N
l | City name | The name of the city where the Authorised representative is located. | Y | Y | Y | Y
m | PostcodePostal code | The postal or zip code where the Authorised representative is located | Y | Y | Y | Y
1.3.4 | Submitter’s details if subcontactor of manufacturer or authorised representative
a | Registered commercial name of company | The name of the company (the legal entity) for this trend report adverse incident
b | Contact's first name | First name of the person to contact about the trend report incident
c | Contact's last name | Last name of the person to contact about the trend report incident
d | E-mail | The email address for the submitter
e | Phone | The telephone number for the company
f | Country | The country where the company is located.
g | Street | The street of the submitter
h | Street number | The number of the manufacturer on the street
i | Address complement | The address where the company is located- e.g. building name
j | PO Box | The PO box of the submitter
k | City name | The name of the city where the company is located.
l | Postal code | The postal or zip code where the company is located
2. Medical Device Information
2.1 | Device scope type in trend
a |  | Select the device level by which the MTR is prepared scope of the devices in the trend report. Only one selection is possible. MDR IVDR and MDD, AIMDD legacy devices could be reported by EMDN code category, group or type of devices, Basic UDI-DI(s)/Eudamed DI(s), individual device or individual UDI-PI(s)/LOT number. For the IVDR and IVDD the reporting could be based on Basic UDI-DI, individual device(s) and UDI-PI/LOT number(s). The old devices could be reported only by device(s) name(s) and reference number(s).The old devices the reporting could be reported based only on device(s) name or names and reporting by the category of devices is not possible. | Y | Y | Y | Y
2.2 | Device(s) scope Identification | [5 empty] | Based on Eudamed rules one of the filds is mandatory
a | EMDN code(s), Nomenclature text for each code | The field is used when the device scope on the MTR is Category/Group/Type from Medical device nomenclature (EMDN) i.e. the scope is on higher level than the devices are registered in Eudamed. Possible only for MDR, IVDR and legacy devices. The data should be provided manually. Used in cases the group of devices included into the MTR is on higher level than registered into the Eudamed ore one UDI-DI includes severals EMDN codes | Y | Y | Y | Y | One of either a,b,c or d
b | Basic UDI-DI(s)/Issuing entity / Eudamed DI(s) | The field is used when the device scope on the MTR is on Basic UDI-DI(s)/Eudamed DI(s) level. Possible only for MDR, IVDR and legacy devices. The data stored in Eudamed will be autopopulated. Text for Issuing entity is missing | Y | Y | Y | Y
c | UDI-DI(s)/EUDAMED ID(s),Trade name(s), Catalogue/Ref number(s) | The field is used when the device scope on the MTR is on device level.The data stored in Eudamed will be autopopulated and the data related to the old devices it should be filledprovided in manually | Y | Y | Y | Y
ed | UDI-PI(s)/Lot/Batch number(s), Trade name(s), Catalogue/Ref number(s) | The field is used when the scope of the MTR is on the batch level. In case the MTR scope is production identifier and more than 5 batches are affected, the manufacturer should upload a list which identifies the devices and production identifiers (UDI-PI, (LOT/Batch number) The data stored in Eudamed will be autopopulated and for the old devices it should be filled in manually | Y | Y | Y | Y
2.3 | Risk class of device(s) when placed on market | Data content according to the specifications of Eudamed
a | MDD/AIMDD/IVDD Class | The data stored in Eudamed will be autopopulated. Indicate the relevant Device Class for the medical device to which this report refers. If unknown please select risk class to the best of your knowledge and update on follow up/final. Eudamed will provide the information in casest the device is registered into Eudamed. | y | y | y | y
b | MDR/IVDR Class | The data stored in Eudamed will be autopopulated. MDR and IVDR fields set up as according to proposal in Eudamed The information will be attached to the UDI (Basic UDI-DI) in the UDI database, therefore it will come automatically with the Basic UDI-DI (except for custom-made device which will not have a Basic UDI-DI). If unknown please select risk class to the best of your knowledge and update on follow up/final | y | y | y | y
2.4 | Notified Body Information
a | Notified body (NB) ID number(s) (if applicable) | Autopoulated by Eudamed or manually filled in (based on the information stored in the Eudamed) . All the NB(s) listed will have full access to the Trend report | y | y | y | y
3. | Description of Trend
a | Background information on the trend, including probability of problem arising and the predicted risk to users. | Provide the following information: description of the incidents, the expected side effects or expected erroneous results identified. The methodology used to identify the detected trend, the established and identified threshold, probability of problem arising and the predicted risk to users. | y | y | y | y
3.1 | Coded information on trend
3.1.1 | Medical device problem information
a | IMDRF Medical device problem codes (Annex A)/in-house codes | Please enter the most relevant lowest level observation as Choice 1. You may chose up to 6 other different codes that describe the medical device problem. In case you can't find a level 3 observation, but a suitable level 2 code, then please use this code as Choice 1 and explain briefly in the text box why no level 3 code was chosen. You can propose a new IMDRF code/term. If no code on level 2 can be chosen, select one from level 1 and propose a new IMDRF code/term for levels 2 and 3. The aim is always to code to the most appropriate level. If no code (on any hierarchy level) can be found briefly explain why. (This is a way to propose new IMDRF terms which could be incorporated in the nomenclature during a maintenance session). The in-house codes are allowed to use within the transitional period. When using the in-house codes please provide the verbal text description of the code. | y | y | y | y
3.1.2 | Healt effect
a | IMDRF 'Clinical signs, symptoms, and conditions’ and (annex E) and health effect (annex F) terms and codes/in-house/MedDRA codes. IMDRF 'Health Effect' terms and codes (Annex E, F) or in-house/MedDRA codes | Please enter the most relevant IMDRF "Clinical signs, symptoms, and conditions codes" (Annex E) and IMDRF "Health impact" codes (Annex F). You can choose up to 6 codes to describe: 1) clinical signs, symptoms, conditions 2) health impact. Please enter the most relevant lowest level observation as Choice 1. You may chose up to 2 other different codes that describe the health effect The aim is always to code to the most appropriate level. The in-house/MedDRA codes are allowed to use within the transitional period. When using the in-house or MedDRA codes please provide the verbal text description of the code. | Y | Y | Y | Y
3.2 | Root Cause Analysis
a | Summary of the root cause analysis | Concerns only the final report. When a root cause is identified the MF reporter enters its final comments on the root cause and conclusion. The investigation process is described detailed on the trend report document. | N | N | N | Y
3.32 | Actions to be taken or implemeted to reduce risks for the users/patients
a | Actions to be taken or already implemented to the reduce the risks for the users/patients Corrective actions information and justification if the corrective or preventive actions has not been taken | Provide an overview about the CAPA or FSCA or other measures decided to adopt in order to reduce the severity and/or the increase of frequency of the reported trend. The corrective and preventive action are described detailed on the trend report document and the FSCAs are follwed through Eudamed. | N | N | N | Y
b | FSCA reference number | Provide the reference number of the initiated FSCA | N | N | N | N
3.4 | General comments
| Free text field |  | N | N | N | N
3.5 | Upload the Trend report document (pdf)
a | Trend report document to upload (pdf) | The form registers the MTR into Eudamed and provides initial information of the detected trend. The detailed information about the trend and its investigations is provided on the Trend Report document. Firstly, fill in the form, secondly upload the document and thirdly submit it to Eudamed. | Y | N | Y | Y
Note on C67: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I removed the added text since the reporting is nbot based on the device level.