Questions and Answers on Trend reporting as outlined in the Regulation (EU) 2017/745 on medical devices (MDR) and the Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR)

Introduction

This document aims to explain and clarify questions related to Trend reporting as outlined in Article 88 of the Regulation (EU) 2017/745 on medical devices (MDR) and Article 83 of Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR). A common understanding related to Trend reporting is deemed essential for an effective and harmonised implementation of the vigilance requirements under the MDR and IVDR.  To that end, this document is intended as a supporting tool for use by Competent Authorities, Economic Operators and other relevant parties to be read, in conjunction with the MDR and IVDR, to provide a further clarification of the MDR and IVDR vigilance requirements and obligations.

The Trend Reporting requirements, as defined in the Article 88 of the MDR and in the Article 83 of IVDR, are respectively applicable from the 26th May 2021 and 26th May 2022.

The Trend Reporting under the MDR requires a consistent, standardized and systematic review of all incidents and expected undesirable side effects apart from serious incidents, by medical device manufacturers. As far as, IVDR requirements are concerned, a standardized and systematic review of all incidents and expected erroneous results, apart from serious incidents, is needed.

Undesirable side-effects must be documented and considered in the manufacturers risk management file, trending procedures and the benefit risk assessment of the device in order to detect potential trends that must be reported under Article 88 of MDR. If the manufacturer is uncertain about whether an incident is an expected undesirable side-effect and whether the incident falls under the definition of a serious incident, it must nevertheless submit a serious incident report within the reporting timelines. As far as, IVDR requirements are concerned, a standardized and systematic review of all incidents and expected erroneous results, apart from serious incidents, is needed.

The MDR trend reporting is applied to the MDR and legacy devices. Similarly, the IVDR trend reporting is applicable to the IVDR and legacy devices. For the requirements on trend reporting of old devices please refer to Question n.4.

The Terms and Concepts, useful for Trend Reporting, as outlined in Section 2 of Chapter VII of the Regulation (EU) 2017/745 on medical devices, are defined on the document Questions and Answers on Vigilance Terms and Concepts (add reference when available).

For an easier reading, consider that the concepts defined in this document are applicable to the corresponding articles of the Regulation (EU) 2017/746 on in vitro diagnostic medical devices (IVDR).

- What shall a manufacturer report with a Trend Report?

Article 88 (1) MDR and Article 83 (1) IVDR outline the requirements for trend reporting.

The Competent Authority shall be notified if a statistically significant increase of incidents or expected undesirable side effects (MDR) or expected erroneous result (IVDR) is detected and that change could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. Trends should be identified by the manufacturer as they can be indicative for a change in the risk-benefit ratio.

- When does a manufacturer have to report incidents, expected undesirable side effects or expected erroneous results to a Competent Authority using a Trend Report?

The manufacturer has to record incidents and expected undesirable side effects in accordance with Articles 83-86 of the MDR and incidents and expected erroneous results in accordance with Articles 78-81 of the IVDR prescriptions for the post market surveillance.

As specified in the Article 83 (2) of the MDR and Article 78 (2) of the IVDR, a manufacturer in its post market surveillance system has  to actively and systematically gather, record and analyse, any relevant data on the quality, performance and safety of a device throughout its entire lifetime, in order to be able to draw the necessary conclusions and to determining, implementing and monitoring any preventive and corrective actions. As also specified in Article 83 (3) of the MDR and Article 78 (3) of the IVDR, data gathered by the manufacturer's post-market surveillance system shall, among other things, be used to detect and report trends in accordance with Article 88 of the MDR and Article 83 of the IVDR.

Whenever a statistically significant increase in the frequency or severity of non-serious incidents or of expected undesirable side effects (MDR) or expected erroneous result (IVDR) is detected and results in a change of the risk evaluation and has a significant impact on the benefit-risk analysis, the manufacturer shall draw up a trend report as defined in the Article 88 of the MDR and in the Article 83 of the IVDR.

It is worth mentioning that while assessing the increase in the events frequency and severity, it is implied that the risk management system must be based on quantitative categorization of probability of occurrence of harm and severity of harm versus the option to use either qualitative or quantitative categorization as required by EN ISO 14971: Medical devices - Application of risk management to medical devices.

A trend report can be drawn up when the manufacturer has performed a monitoring of incidents over time and has compared the severity and the frequency of the occurrences with the threshold defined during the design process and documented in the risk analysis.

- How shall the manufacturer report the “expected undesirable side effects” in trend reporting?

Expected undesirable side effects must be clearly documented in the product information and quantified in the manufacturer’s technical documentation. They must also be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal condition of use (Section 8 of Annex I of MDR).

As specified in Article 88 of the MDR, expected undesirable side-effects should be reported through trend reports. To meet the definition of an expected undesirable side-effect as referred to in Question 8 of the Q&A document on Vigilance terms and concepts, the manufacturer should demonstrate, within the defined timeframe outlined in Article 87 (3) to (5) MDR, that the event is not linked to any device malfunction nor any deterioration in the characteristics or performance of the device nor to an inadequacy in the information supplied by the manufacturer.

If the manufacturer fails to do so, as according to Art. 2(64) MDR all undesirable side effects are considered incidents by definition and the incidents should not be considered an expected undesirable side effect. It has to be evaluated according to Article 87 of the MDR and if the incident has led, might lead or might have led to a serious deterioration of the patient’s state of health, it has to submit an individual Manufacturer Incident Report (MIR).

- What about old device in the Trend Report process and how to submit them?

For the device types which have been placed on the market only prior to the date of application of the respective Regulations, so called “old devices”, the MDR and IVDR requirements related to trend reporting are in principle not applicable. MDR/IVDR requirements that impact the device documentation (the labelling, the technical documentation to be drawn up) or the conditions for the placing on the market of the device, do not apply to old devices.

If possible, it is recommended to follow the same procedure for trending with the old devices as for the MDR/IVDR devices.  If the manufacturer does not apply the same process for the old devices, then the manufacturer is recommended to inform the Competent Authority about the submission process.

When a manufacturer identifies a trend regarding an old device, it is possible to have 2 different situations:

- When the legacy or MDR/IVDR device, which is covered by the scope of the trend report, is similar to the corresponding old device (e.g implant with same or similar design), the events linked to the old device are recommended to be included in the MDR and IVDR trend report.

  - In cases the old devices are not included into the MDR or IVDR trend report the manufacturer could prepare an additional trend report covering the old devices.

- In cases the trend report covers only the old devices, the manufacturer should submit its trend report to the national vigilance system or to Eudamed if requested by the Member State to which it is addressed.

When Eudamed trend report interface will be available, it can be used as for trend reports that cover only old devices as for trend reports that covers legacy and MDR/IVDR devices, according to Question 15 of this document. In the meantime, the old trend report form can be used.

- What happens if a NEW unexpected undesirable side effect/erroneous result or incident is identified?

Unexpected undesirable side-effects and erroneous results or newly identified incidents are events that were not considered or addressed by the manufacturer in its risk analysis and should be:

- reported in accordance with Article 87/ 83 the MDR/IVDR (in case of serious incidents) and

- documented and analyzed during the update of the risk analysis and the risk management file.

The update of the risk-benefit assessment should take the identified new risks in consideration and labeling updates implemented as required.

Whenever a statistically significant increase in the frequency or severity of incidents or of unexpected undesirable side effects/erroneous result is identified and confirmed according to the manufacturer internal procedures, the manufacturer should conduct a full assessment on impact of the trend on the benefit-risk ratio of the product and shall draw up a trend report as defined in the Article 88/83 of the MDR/IVDR.

- How is the PMS plan linked with the trend reporting process?

Taking into account the requirement of Post Market Surveillance (PMS) (MDR and IVDR, Chapter VII: Post-market surveillance, vigilance, and market surveillance) manufacturers do not only need to perform trend reporting as per Articles 88 for MDR and Article 83 for IVDR, but they must also include in the PMS plan the methods and protocols to manage the events subject to trend reporting. This includes the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period (per Annex III, Section 1.1 b: Technical Documentation on Post-Market Surveillance).

As per Article 88 (1) (MDR) and Article 83 (1) (IVDR), the manufacturer shall specify, in the post-market surveillance plan referred to in Article 84 for MDR and Article 79 for IVDR:

-  how to manage the incidents subject to the trend report;

-  the methodology used for determining any statistically significant increase in the frequency or severity of such incidents (or change in performance considering the IVDR);

- the observation period.

Article 84 (MDR) and Article 83 (IVDR) specify that the requirements for the post-market surveillance plan and are set out in Section 1 of Annex III which says that “methods and protocols to manage the incidents subject to the trend report as provided for in Article 88 (83), including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period.”

For more detail about the post market surveillance please refer also to “PMS Guidance” (add a link when it is available).

- How to link the trend reporting process to general Quality Management System?

The general requirements on the quality management system (QMS) of medical devices and IVD’s are described in section 4.1 of the ISO 13485 as well as in Annex IX and the Article 10 of the MDR /IVDR. The Implementation of the quality management system shall ensure compliance with this Regulation and the applicable standards.

The trending process is one of the key elements of vigilance, risk management, PMS processes and of the QMS. Manufacturers of medical devices and IVD’s shall plan and implement the monitoring, measurement, analysis needed to demonstrate the conformity, the performance, and the safety of their products and to ensure the effectiveness of the quality management system.

Therefore, according to  Article 10 (13) of the MDR and Article 10 (12) of the IVDR, manufacturers shall have a system for recording and reporting of incidents and field safety corrective actions as described in Articles 87 and 88 of the MDR and Articles 82 and 83 of the IVDR.

The risk management , the clinical evaluation, and PMS processunder which manufacturer operates are essential key elements in trend reporting being the basis of determining:

- when trends become significant,

- when the acceptable (lowest) risk level (severity-harm combination) might become threatened,

- when the risk-benefit analysis of residual risk has led or may lead to unacceptable risks to health or safety of patients,

- when the residual risk does not constitute an acceptable risk when weighed against the benefits to patient,

- when actions are necessary by manufacturers to address unacceptable risks.

- Which incidents/undesirable side-effects/expected erroneous results from a geographic perspective have to be included in Trend Reporting?

The manufacturer has to record incidents and expected undesirable side effects/expected erroneous results in accordance with Articles 83-86 of the MDR and Articles 78-81 of the IVDR prescriptions for the post market surveillance.

In order to have statistically significant data as a basis for a trend report, the manufacturer has to take in account all the incidents related to a specific trend and from a geographic perspective has to consider all the incidents that have occurred worldwide. Data collected worldwide by the manufacturer during PMS activity, are all useful to define the statistically significant dataset which establish a trend. This trend will be notified to the countries in which MDR/IVDR is applicable and where related incidents/expected undesirable side effects/expected erroneous results occurred.

Therefore, the detection of trends is based on the data collected on incidents occurring in Europe and worldwide.

Considering limitation to those countries where incident reporting is applied. the manufacturer should prepare a trend report based on the MDR/IVDR when the trend includes at least one event7 that occurred in the European Economic Area (EEA)+ Türkiye (TR)+ Northern Ireland (XI).

- What if new events occur in countries that were not included in the initial report?

If during the trend report cycle the new events6 are detected in countries (EEA countries, TR or XI) which were not included in the initial Manufacturer Trend Report (MTR) ,the manufacturer has to add the relevant CA’s at the latest at the time of the final report or as applicable in case of a follow up report.

- Can a trend report include only one device or potentially also a group of devices?

As indicated in the Article 88 (1) of the MDR, a trend report could be performed for a single device as well as a category or group of devices if they are impacted by the same incidents/ undesirable side-effects.  The manufacturer can select the device scope by choosing between the medical device category (the first hierarchy level on EMDN) or group (the second hierarchical level on EMDN) , Basic UDI-DI(s)/Eudamed DI(s), UDI-DI(s)/Eudamed ID(s) and UDI-PI(s)/Lot/batch number(s). The old devices could be reported only by the device name and reference number.

Article 83 of the IVDR does not specify that a trend report could be applicable to a category or group of products. The manufacturer can select the device scope covered in one MTR by choosing between Basic UDI-DI/Eudamed DI, UDI-DI/Eudamed ID and UDI-PI(s)/Lot/batch number(s) and the old devices by the device name and reference number.

In case the MTR scope is the production identifier and the MTR cover more than 5 batches the manufacturer should add a list which identifies the devices and the production identifiers (UDI-PI, (LOT/Batch number). The list should be .csv format and it will be uploaded  into Eudamed. If the affecter batches are 5 or lees the data could be entered directly to Eudamed interface.

For this reason, in the Section 2.1 of the Trend Report Form the manufacturer can select the device scope type choosing between the medical device high level EMDN code i.e. group of medical devices, Basic UDI-DI/Eudamed DI, UDI-DI/Eudamed ID and UDI-PI/Eudamed ID (lot/batch number). The old devices could be reported only by the level of device name(s) and reference number(s).

Only one choice is possible in the device scope type.

- Can the device scope and the scope of the trend change during the trend reporting process and the trend report life cycle?

When a trend report is drawn up and a device scope is defined it has to be kept the same during the trend reporting process and trend report life cycle. This is because a trend report is based on statistical data collected by the manufacturer with a specific methodology and they have to be referred to a specific device scope in order to avoid any confusion and to allow the Evaluating Competent Authority to carry out the appropriate assessment on the data provided.

If a new device scope is identified for the same device/category/group of devices, it is necessary to create a new trend report with its own process and life cycle.

If during the life cycle of the trend report, after further investigations, the manufacturer changes the device scope, it could be included in the trend report follow up ONLY if the assessment, which led to the trend report, remains unchanged (for example a new batch of device is identified).

- If a device has different problems, how should they be reported?

If a device has different problems that meet Article 88 of the MDR or Article 83 of the IVDR requirements, each problem should be reported in a separate trend report.

- With reference to Article 88 of the MDR and Article 83 of the IVDR and the criteria for trend reporting, what is meant by a ‘significant impact on the benefit-risk analysis’?

A ‘significant impact on the benefit-risk analysis’ is whenever the increase in the frequency or severity of incidents, that are not categorized as serious incidents or that are expected undesirable side effects, results in a change of the risk evaluation (EN ISO 14971). A significant increase in the severity could be defined as possible consequence of the incident and / or increase of the probability of occurrence of harm in such a way that could lead to unacceptable Residual Risk and significant impact on the benefit-risk analysis. An example of this can be, the change of a risk zone from low to medium, or increase from medium to high, within the risk matrix as suggested in accordance with EN ISO 14971.

In line with Article 88 of MDR and Article 83 of IVDR and in reference to Sections 1 and 5 of Annex I / Chapter I, the final decision to report is complemented by a health hazard assessment in which the residual risk is weighed against the intended benefits.

- What is the “threshold” mentioned in the section 3 (b) of the Trend Report Form?

The established threshold value for reporting refers to the threshold defined by the manufacturer in its risk assessment in a way to avoid over or under-reporting. In this section the manufacturer should describe how he has determined potential risk and which signal was detected.

Thresholds of risk for trend reporting should be defined at the beginning of the designing and manufacturing processes in accordance with the risk-benefit analysis and reference to them should be reported in trend report.

With particular reference to medical devices, risk can be considered as the results of two components:

  - probability of harm occurring;

  - consequences of the harm, meaning severity of the damage.

Benefit is defined as the benefit of a medical device which is related to the likelihood and extent of the health improvement expected from its use.

With regard to in-vitro diagnostic devices, risk can be considered as the result of three components:

  - the expected performance during its clinical use;

  - the clinical outcome expected from this service;

  - factors relevant to the risks and benefits of other treatment options.

As the standards EN ISO 14791 do not define exactly the “acceptable risk” the evaluation is required by the manufacturer who should have a suitable system in place for proactive scrutiny.

Some qualitative methods can be identified in agreement with the EN ISO 14971.

Examples of matrix for qualitative and semi-qualitative risk assessment can be identified consulting the EN ISO 14971.

- Which Trend Report Form has to be used before and after Eudamed becomes fully functional?

Trend reporting via the electronic system (Eudamed) referred to Article 92 of the MDR and Article 87 of the IVDR using the Trend Report interface will be applicable only when Eudamed becomes fully functional, and becomes mandatory 6 months after publication of the notice in the Official Journal of the European Union.

As per Article 92 (6) of the MDR and Article 87 (6) of the IVDR the trend reports shall be automatically transmitted upon receipt via the electronic system to the Competent Authorities of the Member States in which incidents occurred.

Eudamed will send a notification of the trend report also to the National Competent Authority (NCA) where the manufacturer, or its authorized representative, has his registered place of business.

In the meantime, until Eudamed is fully functional, the MDR/IVDR Trend report requirements are applicable and the alternative administrative technical solutions have to be adopted as per guidance MDCG 2021-1 and MDCG 2022-12. The current Trend report form should be used and it should be sent to the national vigilance systems of Member States.

- Which are the documents related to the Trend Report when Eudamed is fully functional?

When the manufacturer issues a Trend Report, he has to draw up 2 documents:

- Eudamed Trend Report interface. It is the Manufacturer Trend Report (MTR) form. The data will be  directly entered in Eudamed including the following information: Member states where the incidents occurred, the administrative information, description of the devices and the detected trend (background, established threshold, methodology, observation period, and actions to be taken or already implemented to the reduce the risks for the users/patients).

- Trend Report document. The trend report document.- includes additional information about the affected devices, the methods used to detect the trend, description of the detected trend, its investigations, conclusions made and information related to the corrective and preventive actions.The trend report document should be uploaded into Eudamed on the PDF-form. The Trend report document is not mandatory with the initial trend report, but it is expected with the follow- up and final reports.

The need for a separate Trend Report document is related to the reduced amount of information that can be introduced in the Minimum Viable Product version of Eudamed. As mentioned above, the Eudamed Trend Report Interface (and its related helptext) will be applicable only when Eudamed becomes fully functional.

- Can a Trend Report lead to the definition of a corrective action/preventive action?

During the assessment of the trend report data, the manufacturer can issue a corrective action including field safety corrective action/preventive action to reduce or prevent risk for patients and users of the device. In this context the manufacturer should notify a Final trend report and initiate an appropriate action (e.g. FSCA, CAPA, update of labeling, training).

As specified in the Regulation (Annex I MDR 2017/745 and IVDR 2017/746), a manufacturer in its own risk management system has to implement all actions to reduce the risk.

If a manufacturer does not implement any action or any measure in order to reduce the risk he should provide a substantiated justification or a rationale about it to the CA. This justification must be included in the Trend report interface (see Q.15).

If the evaluating CA is not satisfied with the action taken by the manufacturer, it can require additional corrective actions as appropriate, according to Article 88 (2) of the MDR and Article 83 (2) of the IVDR.

- What types of request can a Competent Authority make after the assessment of a trend report?

The Competent Authorities, as specified in Article 88 (2) of the MDR and Article 83 (2) of the IVDR

, may conduct their own assessments on the trend reports and require the manufacturer to adopt appropriate measures in accordance with the MDR/IVDR in order to ensure the protection of public health and patient safety.

When a trend report is received the CA assesses it. This assessment could also lead to a request of re-evaluation by the manufacturer.

The appropriate measures that a CA may request to a manufacturer can be of various types, as long as all those included in the Regulation itself. For example, it can be asked to issue a corrective/preventive action or any other kind of measure in order to ensure the protection of public health and patient safety.

- How can the manufacturer manage the Custom-Made Device (CMD) in the scope of Trend Report?

The same process related to the identification of a Trend applies also on the CMDs.

With the absence of stated exceptions, CMD manufacturers must meet nearly all of the MDR requirements.

As defined also in the guidance MDCG 2021-3, for risk management, post-market surveillance and clinical evaluation life cycle processes as defined by the MDR, CMD manufacturers should apply these obligations to groups of devices with the same intended purpose, materials used, process utilized, same principal design etc. and not to each individual CMD.

As a consequence, the requirements of Article 88 of the MDR apply to CMD and also the registration to the Eudamed according to the guidance MDCG 2021-13 rev.1.

- What are some examples for reporting a Trend Report?

In order to further clarify what shall be reported in a trend report, some examples are provided below:

Examples for MD:

- A trend of expected undesirable side effects causes a risk zone changes from low to medium, or an increase from medium to high, within the risk matrix as suggested in accordance with EN ISO 14971, Medical devices, Application of risk management to medical devices;

- A number of similar incidents, according to Article 2 (64) of the MDR, that are not serious incidents reported by different healthcare facilities shows a statistical increase with a daily or monthly frequency;

- An increase in an incident scenario that does not meet reporting requirements of Article 87 of the MDR (harm severity: minor) causes a risk management threshold breach and leads to an increase in probability from “remote” to “occasional”

- skin irritation as a result of Medical Devices applied with direct dermal contact.

- Abdominal pain and sickness in the first few weeks after implantation of a contraceptive coil

- Paradoxical Hyperplasia after using a cryolipolysis device.

Examples for IVD:

- The change in upper and lower detection limits.

MDCG 2021-25 Regulation (EU) 2017/745 - application of MDR requirements to ‘legacy devices’ and to devices placed on the market prior to 26 May 2021 in accordance with Directives 90/385/EEC or 93/42/EEC

MDCG 2022-8 Regulation (EU) 2017/746 - application of IVDR requirements to ‘legacy devices’ and to devices placed on the market prior to 26 May 2022 in accordance with Directive 98/79/EC

referred to in Sections 1 and 5 of Annex I MDR and IVDR

Side-effects cannot be the result of a malfunction, deterioration in the device’s characteristics or performance, or an inadequacy in the information supplied by the manufacturer.

Reference of the published Q&A document on Vigilance terms and concepts (ADD A REFERENCE).

The term “events “ in this context covers incidents, undesirable side-effects and expected erroneous results

EN ISO 14971 Medical devices, Application of risk management to medical devices.

Probability of occurrence of harm = the probability of hazard or hazardous situation multiplied by the probability of hazard or hazardous situation causing harm (Questions and answers on vigilance terms and concepts as outlined in the Regulation (EU) 2017/745 on medical devices)

Residual Risk = risk remaining after risk control measures have been implemented (EN ISO 14971 definition).

For the IVDs Clinical benefit consult also the MDCG 2022-2- Guidance on general principles of clinical evidence for In Vitro Diagnostic medical devices (IVDs). “It should be recognized that the concept of clinical benefit for in vitro diagnostic medical devices is fundamentally different from that which applies in the case of pharmaceuticals or of therapeutic medical devices, since the benefit of in vitro diagnostic medical devices lies in providing accurate medical information on patients, where appropriate, assessed against medical information obtained through the use of other diagnostic options and technologies, whereas the final clinical outcome for the patient is dependent on further diagnostic and/or therapeutic options which could be available.”

MDCG 2021-1 “Guidance on harmonized administrative practices and alternative technical solutions until Eudamed is fully functional”

MDCG 2022-12 Guidance on harmonized administrative practices and alternative technical solutions until Eudamed is fully functional (for Regulation (EU) 2017/746 on in vitro diagnostic medical devices)

Refer also to MDCG 2021 – 3 guidance “Questions and Answers on Custom-Made Devices” Question 8 and 9.

Question 8 MDCG 2021 -3: 8. What are the obligations of CMD manufacturers according to the MDR?

.

2021-13 rev.1 Questions and answers on obligations and related rules for the registration in EUDAMED of actors other than manufacturers, authorized representatives and importers subject to the obligations of Article 31 MDR and

Article 28 IVDR