Medical Devices

MDCG 2024-XX

Guidance on post-market surveillance of medical devices and in vitro diagnostic medical devices

Month 20xx

Medical Device Coordination Group Document                                    MDCG 2024- xx

This document will be has been endorsed by the Medical Device Coordination Group (MDCG) established by Article 103 of Regulation (EU) 2017/745. The MDCG is composed of representatives of all Member States and it is chaired by a representative of the European Commission. The document is not a European Commission document, and it cannot be regarded as reflecting the official position of the European Commission. Any views expressed in this document are not legally binding and only the Court of Justice of the European Union can give binding interpretations of Union law.

Table of contents

i.	Introduction	3

ii.	Scope and Objective	3

iii.	Terms and Definitions	3

1 The PMS System required by the MDR/IVDR	3

1.1 The PMS plan	4

1.2 PMS for custom-made devices (MDR only)	7

2 The main activities within a PMS system	7

2.1 Selecting appropriate sources for PMS	7

2.2 	Collecting necessary data	9

2.3 Assessment and analysis of data	9

2.4	Drawing conclusions and determining the need to take action	9

2.5	Information on Preventive or Corrective Actions (CAPA) according to Article 83 (4) MDR and Article 78 (4)	10

3 The interactions of the PMS processes with other key processes of the QMS	10

3.2 Update the design and manufacturing information, instructions for use and labelling  - Art 83(3b) MDR /Art. 78 (3b) IVDR	10

3.3 Update clinical evaluation / performance evaluation	11

3.4 Update the summary of safety and (clinical) performance (SS(C)P)	11

3.5 Identify needs for preventive, corrective, or field safety corrective actions	12

3.6 Identify options to improve usability, performance, and safety of the device	12

3.7 Contribute to PMS of other devices	12

3.8 Detect and report trends	12

References	12

Annex I Overview of PMS obligations in MDR and IVDR	14

Annex II Scenarios MDR and IVDR how to use PMS data to update other processes	16

# Introduction

A fundamental goal of the Medical Device Regulation (EU) 2017/745 (MDR) and the In Vitro Diagnostic Medical Device Regulation (EU) 2017/746 (IVDR) is to establish a regulatory framework for medical devices which ensures a high level of safety and health, thus protecting the health and safety of patients, users and other individuals, while supporting innovation.

A device may only be placed on the Union market if it meets the requirements of Union legislation, as defined by the MDR and the IVDR. Article 10 of the MDR/IVDR lays down the general obligations of medical device manufacturers and requires manufacturers to establish, document, implement, maintain, keep up to date and continually improve a Quality Management System (QMS).

Once placed on the market, European regulations oblige the manufacturer to perform post-market surveillance (PMS) of the device throughout its lifetime

# Scope and Objective

Unless otherwise stated, this guidance is applicable to all medical devices and in vitro diagnostic devices (IVDs), without prejudice to national legislation. PMS requirements also apply to “legacy devices” per Article 120(3) of the MDR and Article 110(3) of the IVDR.

The objectives of this guidance are:

- To describe the PMS system and processes required by Article 83 and 84 MDR/ Article 78 and 79 IVDR;

- To describe the main elements and activities within the PMS system.

- To clarify the interactions of the PMS processes with other key processes of the quality management system (QMS);

# Terms and Definitions

For the purpose of this document, the definitions of Article 2 of the MDR and IVDR apply. For terms and definitions that are not provided in the MDR and IVDR, but that are explained in existing MDCG guidance documents, we will reference these in the footnotes.
When using the term ‘device’ in this guidance document it will be understood to include medical devices, accessories to medical devices, products listed in Annex XVI of the MDR, IVDs and accessories to IVDs.

# 1 The PMS System required by the MDR/IVDR

Article 10, Section 5(i) of the MDR and 5(h) of the IVDR refers to a PMS system, which per Article 83 MDR/ Article 78 IVDR is to be implemented effectively and in a manner ‘that is proportionate to the risk class and appropriate for the type of device’. The PMS system must be integral to the manufacturer's QMS and suitable for actively and systematically gathering, recording and analysing relevant data on the quality, performance and safety of a device throughout its lifetime and to drawing the necessary conclusions and to determining, implementing and monitoring any preventive and corrective actions. It is in the manufacturer’s interest to use the PMS data to identify options to improve the usability, performance and safety of the device. As part of the QMS, it is required to have an internal reporting system, besides the regulatory reports, to address all relevant QMS processes, e.g., management review.

Planning of PMS starts during the development of a medical device. PMS data collection begins as soon as a manufacturer declares the conformity of a device with the requirements of the applicable regulations and places it onto the market.

The output of PMS is used to continually evaluate and confirm the quality, performance and safety of a device throughout its lifetime and to identify any need for preventive and/or corrective actions.

Per Article 83 MDR/ Article 78 IVDR, manufacturers are required to ‘plan, establish, document, implement, maintain and update’ a PMS system. The PMS system interfaces directly with other processes within the QMS, such as risk management. It therefore requires appropriate structures, procedures, processes and resources for its successful implementation.

Article 84 MDR/ Article 79 IVDR refer to the PMS system being underpinned by a PMS plan, which is part of the technical documentation. See Annex I of this guidance for an overview of PMS obligations laid down in MDR and IVDR.

## 1.1 The PMS plan

There is a requirement to have a PMS plan for each device. The plan specifies the aspects of the device to be monitored, how frequently these will be monitored, and the methods to be applied. Methods are expected to be appropriate considering the risk profile of the device and the rationale for selecting methods needs to be evident.

The requirements of the PMS plan are set out in MDR Annex III Section 1(b) and IVDR Annex III Section 1(b). Of note is that the requirements include the term ‘proactive information collecting process’. Also, MDR Annex XIV Part B and IVDR Annex XIII Part B mention the term ‘proactive’ in the context of Post market Clinical Follow-up (PMPF)/ Post Market Performance Follow-up (PMPF). Use of this term underlines the principle that it is the obligation of the manufacturer to seek available information rather than to rely upon the receipt of information (e.g., complaints data). Data may be proactively collected and/or generated by various means including but not limited to customer surveys, gathering of clinical experience,, feedback from users, screening of scientific literature and of other sources of clinical data, meta-analyses of published clinical data, evaluation of suitable registers and through post-market studies. It is up to the manufacturer to determine appropriate methods considering the rate of innovation and level of research activity within the field.

Although the fourth indent of MDR/IVDR Annex III, Section 1(b) requires that methods and procedures are covered in the PMS plan, in practical terms, it is expected that the PMS plan defines what type of methods are to be applied, while an explanation of how they are to be applied may instead be covered by procedures that are referenced within the PMS plan. It is important to document the roles and responsibilities of various stakeholders within the PMS plan.

The table below provides a summary of the elements to be covered in the PMS plan per MDR/IVDR Annex III Section 1(b):

| Required per Annex III 1(b) | Explanation | Examples of content in the PMS plan |
| --- | --- | --- |
| — a proactive and systematic process to collect any information referred to in point (a). The process shall allow a correct characterisation of the performance of the devices and shall also allow a comparison to be made between the device and similar products available on the market | Manufacturers are expected to collect the information referred to in Annex III Section 1(a) and to use it to address understanding of device usage, safety and performance, usability as well as identify areas for improvement. By collecting information on similar products, the manufacturer can compare their device with similar products on the market (in terms of design or intended use). The information is used to demonstrate that the device remains the state of the art. | - List of activities with references to procedure(s) for collecting the information referred to in Annex III Section 1(a); - Sources of data regarding similar products - Details of the frequency of the activity and the role(s)/function(s) responsible. |
| — effective and appropriate methods and processes to assess the collected data | Methods and processes for assessing the data are expected to be proportionate to the type of device. The methods selected are also dependent upon the type and quality of the collected raw data. It is expected that higher-risk devices will require more specific methods of data collection and analysis. | - Methods to be applied (e.g., qualitative or quantitative, statistical methods) and processes for assessing the collected data, along with rationale; - List of the parameters to be analysed - for example: side effect or type of incident, failure mode, hazard, harm, complaint rate, severity, etc. - List of the reference values - for example: batch, batch quantity, hours/frequency of use, number of devices in use/sold, number of exposures, etc. |
| — suitable indicators and threshold values that shall be used in the continuous reassessment of the benefit-risk analysis and of the risk management as referred to in Section 3 of Annex I | Suitable indicators and threshold values (i.e., limits beyond which action must be taken) are to be established in the pre-market phase. The data obtained through PMS activities is fed back into the risk management process and used to re-evaluate indicators and threshold values. | - Indicators and thresholds relevant for assessing the impact; on hazards and the frequency of occurrence thereof, on estimates of their associated risks, as well as on the overall risk, benefit-risk ratio and risk acceptability. |
| — effective and appropriate methods and tools to investigate complaints and analyse market-related experience collected in the field | Methods and tools are expected to be proportionate to the type of device: For higher-risk devices, more rigorous approaches to investigation and analysis of complaints and other market-related experience may be required. | - Methods to be applied (e.g., qualitative or quantitative, statistical methods) and processes for analysing the collected data, along with rationale; - for example: methods described in IMDRF Annex B, Cause Investigation -Type of Investigation. |
| — methods and protocols to manage the incidents subject to the trend report as provided for in Article 88 MDR/ Article 83 IVDR, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period | Non-serious incidents or clearly documented undesirable side-effects may be subject to trend reporting. It is expected that the methods are defined in the PMS plan and a reference is made to associated procedures. | - References to relevant procedures and statistical methods to be applied along with parameters to be monitored, including threshold values and indicators; - Relevant procedures, documentation and methods/criteria for submitting reports;; - Details of the frequency of the activity and the role(s)/function(s) responsible. |
| -methods and protocols to communicate effectively with competent authorities, notified bodies, economic operators and users | Manufacturers must establish processes to ensure effective communication with competent authorities, notified bodies, other economic operators and users (including patients and customers) with respect to the output of PMS activities. | - References to relevant procedures, i.e., to contact and inform competent authorities, notified bodies, other economic operators and users – including patients and customers; - Details of or reference to forms, communication methods and data transfer tools which may be used |
| — reference to procedures to fulfil the manufacturers obligations laid down in MDR Articles 83,84 and 86 or IVDR Articles 78, 79 and81 | Procedures must be in place to enable the implementation of the PMS system, as well as for generating PMS-related plans and reports. | - References to relevant procedures, i.e., to generate/update PMS plans, reports, etc.; - Details of the frequency of the activity and the role(s)/function(s) responsible. |
| — systematic procedures to identify and initiate appropriate measures including corrective actions; | Procedures are expected to describe triggers for corrective actions and the type of measure to be undertaken (such as documenting a nonconformance, initiating a CAPA, or FSCA) according to the impact or risk posed by the issue identified. | - List of identified triggers which would lead to the initiation of appropriate measures or corrective action. Details of the role(s)/function(s) within the organisation responsible for specific tasks as part of an action plan. |
| — effective tools to trace and identify devices for which corrective actions might be necessary | Tools used to identify and trace devices for which corrective actions might be necessary must be described. Procedures describing controls for non-conforming devices (e.g., recall and quarantine) and identification of relevant economic operators are expected to be referenced. | - List of methods and/or tools to trace and identify devices affected by an issue requiring corrective action; - References to relevant procedures, i.e., to identify and trace devices using specified tools; - Details of the role(s)/function(s) within the organisation responsible for specific tasks as part of a corrective action plan. |
| — a PMCF plan as referred to in Part B of Annex XIV, or a justification as to why a PMCF is not applicable (MDR) — a PMPF plan as referred to in Part B of Annex XIII, or a justification as to why a PMPF is not applicable (IVDR) | The PMCF/PMPF plan is an integral part of the PMS plan. If the manufacturer determines that no specific PMCF/PMPF activities are required, justification is required. | - Description of planned PMCF/PMPF activities and justification; - Reference to a detailed PMCF/PMPF plan/protocol, where appropriate; - Reference to relevant procedures; - Details of the frequency of the activity and the role(s)/function(s) responsible - Justification as to why a PMCF/PMPF is not applicable should be documented within the PMS plan. |

## 1.2 PMS for custom-made devices (MDR only)

The PMS requirements according to Article 83 are applicable to all devices including custom-made devices.

Annex XIII section 5 states that “The manufacturer must review and document experience gained in the postproduction phase, including from PMCF as referred to in Part B of Annex XIV, and implement appropriate means to apply any necessary corrective action. In that context, the manufacturer must report in accordance with Article 87(1) to the competent authorities any serious incidents or field safety corrective actions or both as soon as it learns from them.”

This implies that for a custom-made device the manufacturer needs to have and maintain a PMS system, and documentation regarding the experience gained from the postproduction phase. Articles 84 and 85 of the MDR are not applicable for custom-made devices. Article 86 regarding PSUR is applicable for class IIa, IIb and class III custom-made devices complying with the requirements of the MDR.

# 2 The main activities within a PMS system

The PMS system must be suited to actively and systematically gathering, recording and analysing relevant data on the quality, performance and safety of a device throughout its entire lifetime. This chapter explains the different steps in the cycle of PMS, including the collection, assessment and analysis of post market data and guidance on how to process the outputs. In paragraph 2.1 you can find an explanation of each indent of Annex III part 1(a). It describes the information that must be collected and how this must be utilised.

## 2.1 Selecting appropriate sources for PMS

The manufacturer must decide what sources are relevant and appropriate for the device concerned. When selecting data sources, the different economic operators and users (healthcare professionals, patient or lay users) and the situation in which the devices are used need to be considered. Where examples are given in the table below, these are intended to provide guidance to manufacturer. They should not be understood as mandatory elements. It is up to the manufacturer to select the most appropriate methods, parameters, values, processes and tools depending on the nature and usage of the device. Data quality and integrity must be considered before analysing data to ensure the information is reliable. For example, the use of unverifiable data can lead to over-reaction, as it can be based on non-scientific data sources such as social media and public media. Therefore, appraisal of data and its sources is recommended as part of the PMS plan.

| Required Information as per Annex III 1.a | Examples of how to collect information to satisfy requirements as per Annex III 1.a | Utilise information obtained through PMS |
| --- | --- | --- |
| Information concerning serious incidents, including information from PSURs, and field safety corrective actions | Outputs from procedures to receive, document, investigate and analyse serious incidents received from customers, users, healthcare professionals, patients, other economic operators. Information/feedback gathered during design, implementation and tracking of progress and effectiveness of field safety corrective actions. | Information may contribute to: - identification of emerging risks, or previously unknown side effects; - demonstration of the continued acceptability of the performance and safety of the device. - tracking any emerging or developments to practices or benefits associate with device use (SOTA) - the determination of reportability of incidents to the competent authority/notified body - the identification and investigation of trends - comparison of the device with similar products on the market - identification of new benefits of the device The identification of need for - CAPA - preventative, corrective or field safety corrective action - options to improve the usability, performance and safety of the device, when relevant, to the PMS of other devices - Information to be notified to Competent Authorities/Notified Body: - report serious incidents and corrective and preventative actions - report field safety corrective actions - improvement of internal processes such as internal PMS, vigilance, CAPA, and design control - reporting of trends in accordance with Article 88 - design changes including improving usability, performance and or safety Can improve the - indicators and threshold values for continuous benefit-risk analysis - risk assessment To update - the benefit-risk determination and to improve the risk management - the design and manufacturing information, the IFU and the labelling - the clinical/performance evaluation - the SS(C)P - the technical documentation - the PMS and PMCF/ PMPF plan - risk management including need for CAPA or FSCA - labelling, IFU or other documentation - processes to determine reportability of incidents to the competent authority/notified body To produce - the Post market surveillance report as per Article 85 - the Periodic Safety Update Report as per Article 86 - the PMCF/PMPF evaluation report - clinical/performance evaluation report |
| Records referring to non-serious incidents and data on any undesirable side-effects | Outputs from procedures to receive, document, investigate and analyse: - incidents determined to be non-serious; - complaints and information; - service and repair records, received from customers, users, healthcare professionals, patients, other economic operators. Outputs from procedures to receive, document, investigate and analyse information concerning any undesirable side-effects received from customers, users, healthcare professionals, patients, other economic operators. |  |
| Information from trend reporting | Information obtained via procedures, systems, and methodologies established for trending of incidents and undesirable side effects reported by users and other stakeholders. |  |
| Relevant specialist or technical literature, databases and/or registers | Literature review/ Clinical evaluation process/ External database evaluation process. For example: - Screening literature for reports; - Searching Competent Authority websites and databases, including Eudamed; - Searching device registers or registries, where available. Other published information including clinical guidelines, health technology assessments/reports |  |
| Information, including feedback and complaints, provided by users, distributors and importers; | For example: - Surveys of users, customers, distributors, importers and other stakeholders; - Feedback received during user training, education, workshops; - Reported complaints; Feedback obtained through sales channels including internet or digital sales,and customer service departments or marketing activities. |  |
| Publicly available information about similar devices | For example: - Screening the literature for reports involving similar devices; - Searching Competent Authority websites and databases, including Eudamed; - Searching device registers or registries, where available; - Other published information including manufacturer website, IFU and labelling, SS(C)P; - Information available from third countries (worldwide); - Information on safety and performance gained from conference, congresses and commercial exhibitions. |  |

## 2.2 	Collecting necessary data

The MDR and IVDR require a manufacturer to develop a robust PMS system that enables them to actively and systematically gather, record and analyse relevant data on the quality, performance and safety of a device throughout its entire lifetime. In this way, information is actively and systematically sought to gain data and insight into the real-world performance of the device.

The regulations express an expectation that the PMS plan must include not only methods of data collection but it must be proactive and systematic in the collection of the data. In other words, manufacturers are expected to actively seek to gather information relating to the real-world use of the device. In some instances, manufacturers may be required to generate new data relating to real-world use of the device. This is especially applicable in the case of higher risk devices and/or devices with novel features or applications.

## 2.3 Assessment and analysis of data

The PMS plan must highlight the manufacturers processes and methods that are used to investigate complaints along with gathering information and analysing market related experience that can be collected from the field. The PMS plan must also demonstrate the methods and processes of how the manufacturer will communicate with competent authorities, notified bodies, economic operators, and users.

The PMS plan must also provide a direct reference to the procedures in relation to the output of the PMS data. Outputs include a PMS report for class I devices (Article 85 MDR) and Class A & B devices (Article 80 IVDR) and a PSUR for all other classifications per Article 86 MDR/Article 81 IVDR. The procedures must consider the frequency and thresholds, linked to risk management, that may require an update to the mentioned reports and potentially further updates to the PMS plan.

The PMS plan must also take into consideration other procedures when other appropriate measures may be required such as corrective actions or to identify devices that may have been impacted by such corrective actions. The PMS plan must detail or reference the procedure in relation to these corrective actions (ANNEX III, 1(b) 8th indent).

## Drawing conclusions and determining the need to take action

Articles 83 (3) MDR/ Article 78(3) IVDR state that the data collected by the manufacturer's PMS system must in particular be used for input and to update processes in the system, as explained in the previous paragraphs.

The PMS plan must also provide a direct reference to procedures and frequency related to the output of the plan, namely a PMS report for class I devices per Article 85 MDR/ Class A & B devices per Article 80 IVDR and a PSUR for all other classifications per Article 86 MDR/Article 81 IVDR.

Conclusions on the data collected can have impact in all layers of the manufacturing process, and the chain of economic operators.  Conclusions need to be drawn up in (an update of) a summary report, either the PMS report or the PSUR..Following the outcome of those reports, a review of the current PMS plan may be necessary. This is a manufacturer assessment about the overall device safety and performance, which will drive the level of PMS activity required (based on the device class, historical data, similar products and including the need for PMCF/ PMPF for example).

## Information on Preventive or Corrective Actions (CAPA) according to Article 83 (4) MDR and Article 78 (4)

Preventive Action(s) or Corrective Action(s) which are covered by Article 83(4) first sentence, can be made available to the Competent Authorities either through the PSUR or PMS report. 
All relevant information can be found in the PSUR guidance chapter 2.1.211

# 3 The interactions of the PMS processes with other key processes of the QMS

Aside from establishing, implementing and maintaining a PMS system in accordance with Article 83 MDR/ Article 73 IVDR, the QMS must cover all aspects of a manufacturer's organisation dealing with the quality of processes, procedures and devices.

Article 83 MDR/ Article 78 IVDR specifically requests that information generated from PMS activities must be used as input, on a continuous basis, to other processes within the QMS. In particular, data gathered by the manufacturer's PMS system must be used for the continuous gathering, recording and analysing of relevant data to:

3.1 Update the benefit-risk determination and the risk management.
- Art. 83.3 (a) MDR/ Art. 78.3(a) IVDR

Information obtained through PMS activities must be used in the benefit-risk determination process and in the continuous assessment of the overall benefit-risk profile of the device to ensure that the benefit-risk profile remains acceptable. Furthermore, the data must be utilised to update the risk management documentation.

This means that procedures need to be established to adequately address the interface between PMS and the risk management process. This includes the analysis of the data collected by the planned and established PMS activities and the appraisal of the data. This will determine whether the collected data have any impact on the probability and the severity ratings of existing risks or if new risks were identified. The collected information will inform the overall residual risk and overall benefit-risk profile associated with the device.

The suitable indicators and threshold values used in the continuous reassessment of the benefit-risk analysis of the device must be covered in the PMS Plan and be linked to the risk management documentation (as referred to in Section 3(e) of Annex I).

Changes to the benefit-risk profile of the device as a result of PMS information must be identified, documented and evaluated for acceptability. The need to initiate appropriate measures, including preventive and/or corrective action must be considered.

## 3.2 Update the design and manufacturing information, instructions for use and labelling 
- Art 83(3b) MDR /Art. 78 (3b) IVDR

Information obtained through the PMS process must be used to update the design and manufacturing information, the instructions for use and the labelling. Therefore, a manufacturer must establish procedures to ensure that PMS data is considered as an input for design and further development of the device and provides input for necessary changes to the device, e.g., through identifying previously unknown side-effects and monitoring of the identified side-effects and contraindications or through identifying possible systematic misuse or off-label use of the device. When evidence of new or increased risks is identified, appropriate action must be taken. This could include but is not limited to updates to the labelling and IFU or changes to the design of the device itself.

## 3.3 Update clinical evaluation / performance evaluation

-Art. 61 and 83(3c)MDR/ Art. 56 and 78(3c) IVDR -

Under the MDR, clinically relevant information coming from PMS, in particular the PMCF data, is considered “clinical data”. Annex XIV Part B MDR specifies that PMCF shall be understood as a continuous process that updates the clinical evaluation. Specifically, data obtained by conducting PMCF must be used to:

- confirm the safety and performance of the device throughout its expected lifetime;

- identify previously unknown side-effects and monitor the identified side-effects and contraindications;

- identify and analyse emergent risks on the basis of factual evidence;

- ensure the continued acceptability of benefit-risk ratio;

- identify possible systematic misuse or off-label use of the device, with a view to verify that the intended purpose is correct.

Under the IVDR, performance relevant information coming from PMS, in particular the PMPF data, is considered “relevant scientific data” and “performance evaluation results”. Annex XIII, Part B specifies that PMPF shall be understood as a continuous process that updates the performance evaluation. Specifically, data obtained by conducting PMPF must be used to:

- confirm the safety and performance of the device throughout its expected lifetime;

- identify previously unknown risks or limits to performance and contra-indications;

- identify and analyse emergent risks on the basis of factual evidence;

- ensure the continued acceptability of the clinical evidence and of the benefit-risk ratio.

PMCF/PMPF must be addressed in the manufacturer's PMS plan and must be performed pursuant to a documented method laid down in a PMCF/PMPF plan that specifies the general and specific methods and procedures for proactively collecting and evaluating the collected clinical/performance data. The linkage of the PMS plan to the PMCF/PMPF plan is also described in Annex III, 1(b) indent 10 of the MDR/IVDR. These general and specific methods and procedures of PMCF/PMPF are described in MDCG 2019-9 Rev.1 and MDCG 2022-9.

The outputs of any relevant data associated with PMS must be considered in the process of updating the manufacturer’s clinical evaluation or performance evaluation.

The manufacturer must analyse the findings of the PMCF/PMPF and document the results in a PMCF/ PMPF evaluation report that must be used to update the clinical evaluation/ performance evaluation and the technical documentation. The conclusions of the PMCF/PMPF evaluation report must be taken into account in the clinical evaluation/performance evaluation and the risk management (referred to in section 3 of Annex I). This data must be considered from the perspective of the overall benefit-risk assessment of the device and must be considered in context with existing clinical data or performance data from the device’s previous clinical/performance evaluation. If the need for preventive and/or corrective measures has been identified, the manufacturer must implement them.

## 3.4 Update the summary of safety and (clinical) performance (SS(C)P)

- Article 32 and Art. 83(3d) Article 29 and Art. 78(3d) IVDR

The SS(C)P is applicable to class III and implantable devices (MDR) and Class C and Class D devices (IVDR). The purpose of the SS(C)P is to provide transparency of the clinical data held on the device including aspects of safety and performance to the health care professional and, where applicable, patients. Any new information generated from the PMS system which impacts the benefit-risk-profile (safety/performance profile) of the device may require an update to the SS(C)P. Such updates must be aligned with the information presented in the clinical/performance evaluation report and Periodic Safety Update Report (PSUR). The PMS system must describe how the PMS data is used to update the SS(C)P, when applicable.

## 3.5 Identify needs for preventive, corrective, or field safety corrective actions

- Art. 83(3e) MDR/ Art 78(3e) IVDR

Manufacturers must set up their PMS system to actively not only collect but also analyse and utilise relevant data on the quality, performance and safety of a device throughout its entire lifetime. By this monitoring the manufacturer must draw conclusions whether implementation of preventive and corrective actions is indicated. Furthermore, the manufacturer must set up processes and procedures to integrate the data obtained by the PMS system into the QMS, in order to initiate and document preventive and corrective actions as well as field safety corrective actions to reduce the risk posed by devices already placed on the market.

## 3.6 Identify options to improve usability, performance, and safety of the device

– Art. 83(3f) MDR/ Art. 78(3f) IVDR

As part of the PMS data gathering, manufacturers must establish the processes for receiving, reviewing and evaluating information, including screening of scientific literature, feedback and complaints provided by users, distributors and importers. Data sources may also include correspondence from suppliers, customers, competitors or other third parties. This PMS data must be used to identify options to improve the usability, safety or performance of the device.

## 3.7 Contribute to PMS of other devices

– Art. 83(3g) MDR/ Art. 78(3g) IVDR

The gathered data regarding options to improve usability, performance and safety of one device can contain useful information for the manufacturer’s other devices. Any new or increased risk identified must be considered by the manufacturer as to whether it is likely that other devices that share the same or similar intended purpose, design or process characteristics could be impacted by this new information. When relevant the manufacturer will use the collected information, analysis and conclusions as input to the PMS of other devices. The PMS system in place can utilise PMS data systematically across other devices.

## 3.8 Detect and report trends

- Article 83(3h) MDR and 88 MDR /Article 78(3h) and 83 IVDR)

All information gathered on non-serious incidents, expected undesirable side-effects (MDR) or expected erroneous results (IVDR) must be analysed for any statistically significant increase in the frequency or severity of those events.

Any statistically significant increase in the frequency or severity that could have a significant impact on the benefit-risk profile and which has led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits is subject to trend reporting, according to MDR and IVDR requirements.

# References

The following documents, in whole or in part, have been taken into consideration in the development of this guidance:

- EN ISO 13485:2016, Medical devices — Quality management systems —

- Requirements for regulatory purposes

- EN ISO 14971:2019, Application of risk management to medical devices,

- EN ISO 14155:2020, Clinical investigation of medical devices for human subjects - Good clinical practice

- ISO 20916: 2019 In vitro diagnostic medical devices — Clinical performance studies using specimens from human subjects — Good study practice

- IEC 62366-1:2015, Medical devices — Part 1: Application of usability engineering to medical devices

- CEN ISO/TR 20416:2020, Post-market surveillance for manufacturers

- Guidance for post-market surveillance and market surveillance of medical devices, including in vitro diagnostics, World Health Organization

# Annex I Overview of PMS obligations in MDR and IVDR

| PMS System - MDR Article 83 / IVDR Article 78 |  |
| --- | --- |
| Description | Obligations |
| Comprehensive system to collect experience from use of the devices | MDR/IVDR: Applicable for all device classes. - Proactive and systematic approach to collect data. - Analyse relevant data on the quality, performance, and safety of a device throughout its entire lifetime. - Draw the necessary conclusion. - Connect with other internal processes like CAPA, vigilance, design and labeling processes - Has to be used to update the technical documentation, (except for Custom made devices ) - Integral part of the manufacturer QMS. |
| PMS Plan - MDR Article 84 & Annex III / IVDR 79 & Annex III |  |
| Description | Obligations |
| Describes the implementation of the PMS system to collect data assess the safety and the performance of the devices and the methodology used for collecting and analysing the data. | MDR/IVDR: Applicable for all device classes. - Part of the QMS and the technical documentation. - Define the data to be collected based on the PMS inputs of the manufacturer. - Describes the methods and protocols to monitor trends, to identify significant increase in frequency or severity of incidents - The methods and tools to investigate complaints and analyse market-related experience collected in the field. - The indicators and thresholds to be used to reassess the benefit-risk ratio. - The tools to be used to trace and identify any devices in case corrective actions are needed. - Clinical Evaluation Plan/ Performance Evaluation Plan - The PMCF/ PMPF plan or a justification why PMCF/PMPF is not applicable. - Describes the references to the documented procedures for: - PMS system - Creation of PMS plan - Generation of PMS report or PSUR as applicable. - Processes for corrective actions - Process for risk management - Technical documentation |
| PMS Report, MDR Article 85 / IVDR Article 80 |  |
| Description | Obligations |
| A summary of the results and conclusions of analysis of the PMS data collected | MDR: applicable to class I devices IVDR: applicable to class A and B devices What information must a PMS Report contain at a minimum? - a summary of the results and the conclusions of the analyses/assessment of the post-market surveillance data resulting from the execution of the PMS plan. - A rationale and description of any preventive or corrective actions taken during the period covered by the report. - conclusions on the safety and the performance of the device. The guidance MDCG 2022-2115, although not covering PMSR, may provide useful suggestions on how information can be presented. When does a PMS Report need to be updated? - Update according to the PMS plan or earlier when deemed necessary by the manufacturer. To whom does a PMS Report need to be made available? - To competent authority upon request. |
| PSUR - MDR Article 86 / IVDR Article 81 |  |
| Description | Obligations |
| A summary of the results and conclusions of analysis of the PMS data collected | MDR: applicable to class IIa, IIb and III devices IVDR: applicable to class C and D devices |

# Annex II Scenarios MDR and IVDR how to use PMS data to update other processes

The examples on the next pages are intended to show the possibilities for results following PMS evaluation. It is possible, but not always expected, that every real life scenario would result in all of the actions shown in the examples.

IVD:

Medical Device

MDCG 2022-21 Guidance on Periodic Safety Update Report (PSUR) according to regulation (EU) 2017/745 (MDR), ch 3.2.1.

For a definition and further information on ‘legacy devices’ please see MDCG 2021-25 and MDCG 2022-8

Article 83 (2) MDR/ 78(2) IVDR

These examples are intended to provide guidance to manufacturer. They should not be understood as mandatory elements. It is up to the manufacturer to select the most appropriate methods, parameters, values, processes and tools depending on the nature and usage of the device

similar products is equal to similar devices’ may be understood as devices belonging to the same “generic device group” as defined in Article 2(7) MDR. This should be understood as “a set of devices having the same or similar intended purposes or a commonality of technology allowing them to be classified in a generic manner not reflecting specific characteristics”.

State of the art: Developed stage of current technical capability and/or accepted clinical practice in regard to products, processes and patient management, based on the relevant consolidated findings of science, technology and experience. IMDRF/GRRP WG/N47

PD CEN ISO/TR 20146:2020

See Q&A trend reporting

MDCG 2022-21 Guidance on Periodic Safety Update Report (PSUR) according to regulation (EU) 2017/745 (MDR)

ISO/TR 20416:2020

MDCG 2023-3 Questions and Answers on vigilance terms and concepts as outlined in the Regulation (EU) 2017/745 on medical devices

MDCG 2023-3 Page 4

MDCG 2022-21 Guidance on Periodic Safety Update Report (PSUR) according to Regulation (EU) 2017/745

ISO 9000:2015, Quality management systems, Fundamentals and vocabulary: A process is set of interrelated or interacting activities that use inputs to deliver an intended result. A procedure is specified way to carry out an activity or a process

guidance Q&A on Trend Reporting

MDCG 2019-9 Rev. 1 Summary of safety and clinical performance, MDCG 2022-9 Summary of safety and performance template

MDR Article 2 (67) IVDR Article 2 (70), and ISO 13485

MDCG 2022-21 Guidance on Periodic Safety Update Report (PSUR) according to Regulation (EU) 2017/745