## Blad1
chapter | NCA1 CIE / DA NBOG/ COM / Other | Line number / reference | Type of comment2 | Comment | Proposed change | Subgroup comment3
intro | aesgp | 73 | te | Sentence is incomplete | Change “placing a safe medical device” to “placing a safe medical device on the market” | accept
3 | aesgp | 184 | ed | Confusing phrasing used towards the end of the sentence emphasizing “proactive” PMS. | Change “to actively monitor” to “their active monitoring of” | accept
PMS as an integral part o | aesgp | 220 | ge | What is meant by ‘ threshold values related to certain characteristics’ ? certain reported events? | Change ‘threshold values related to certain characteristics’ to ‘ threshold values related to certain reported events’ or ‘ threshold values defined in risk management file’ | accept text revised
PMS as an integral part o | aesgp | 213 | te | Annex III of the MDR does not have a section 1 (a) | Change reference to 1 (a) to 1.1 (a) | reject - current version is 1(a)
PMS as an integral part o | aesgp | 213 | ed | “(“ at the end of the sentence | Remove the “(“ at the end of the sentence | accept
PMS as an integral part o | aesgp | 216 | ed | ‘ to be made’ is unnecessary | Change ‘ to allow a comparison to be made between’ to ‘ to allow a comparison between’ | accept
PMS as an integral part o | aesgp | 221 | ed | The examples are not necessary | Change ‘ triggers for certain actions (e.g. starting a field safety corrective action, issue a design change etc.)’ to ‘ triggers for CAPA, field safety corrective action, or other relevant actions.’ | CAPA is not wording used in the MDR IVDR is that clear enough for everybody?
4a | aesgp | 262 | te | Threshold values are not necessary in the PMS plan if there is a link to the risk management file that is used as a reference. | Remove ‘threshold values’ from the sentence | partially accepted - text on threshold values and indicators will be reworded to better explain what needs to be laid out in the PMS plan.
4c | aesgp | 319 | te | Annex III of the MDR does not have a section 1 (a) | Change reference to 1 (a) to 1.1 (a) | Reject, current version is 1(a)
4c | aesgp | 330 | ge | Do not understand ‘r/’ at the end of the sentence | Spell out whatever ‘r/’ stands for | accept
4c | aesgp | 296 | ed | Unnecessary “and” at the end of the bullet | Remove “and” at the end of the bullet | accept
4c | aesgp | 316 | ed | Unusual wording with use of laid down | Replace ‘laid down’ with ‘defined’ | Reject - keeping wording of the reulation
4d | aesgp | 343 | ed | ‘only’ not necessary after ‘Class D (IVDR)’ | Remove ‘only’ | accept
4e | aesgp | 356 | ge | Please consider shortening this sentence or separating it into several sentences. It is very wordy and the objective of the sentence is diluted. |  | accept
4f | aesgp | 372 | ed | Missing ‘are’ in sentence | Change ‘ to identify if there any options’ to ‘ to identify if there are any options’ | accept
4f | aesgp | 372 | ed | Make this sentence shorter by including usability with safety or performance | Shorten sentence to say ‘The information should be considered to identify if there are any options to improve the safety, performance, or usability of the device’ | accept
4h | aesgp | 391 | te | There is no mention about the reporting of trends rather only the methods, identification, and analysis of trends | Mention an established procedure for the escalation and reporting of trends. | reject - ref to art 88/83 and guidance for trending
4h | aesgp | 391 | ed | ‘of those events within the quality management system of the manufacturer´ does not add any value to the sentence | Shorten sentence to say ‘ All gathered information on non-serious incidents, expected undesirable side-effects (MDR) or expected erroneous results (IVDR) should be analysed for any statistically significant increase in the frequency or severity’ | accept
5 | aesgp | 403 | ed | 5.1.1 should be referenced | Change 6.1.1 to 5.1.1 | accept
5 | aesgp | 405 | ed | 5.1.2 should be referenced | Change 6.1.2 to 5.1.2 | accept
5.1 | aesgp | 411 | ed | Should specify that the PMS plan is for each device/grouping of devices | Change ‘ should be collected and utilized.’ To ‘ should be collected and utilized for each device or grouping of devices.’ | accept
5.3 | aesgp | 444 | te | The sentence makes it seem as though the PMS plan should focus on the investigation of complaints and not emphasize the other data sources | Change ‘used to investigate complaints along with gathering information and analysing market related experience that can be collected from the field.’ To ‘ used to gather information and analyse market related experience that can be collected from the field. | accept
5.3 | aesgp | 445 | ge | This is a lot to expect out of a plan. Communication with CAs, NBs and EOs should be defined in internal procedures that can be referenced within the PMS plan. | Change to ‘ The post-market surveillance plan should also reference established internal processes of how the manufacturer will communicate with competent authorities, notified bodies, economic operators, and users.’ | accept
5.4 | aesgp | 476 to 485 | ge | We believe this piece could be better organised and the phrasing is confusing. Please consider rephrasing. |  | accept
intro | CA1 | 66 | ed | Typo | “…In-Vitro Diagnostic Medical Devices Regulation.” | accept
intro | CA1 | 58-59 | ed | Since the abbreviations MDR and IVDR have already been defined at the beginning of this document, we suggest that only the abbreviations are used in the following parts of this document for the sake of better readability. Please follow this procedure also for other abbreviations (e.g. line number 172: PMCF/PMPF,…) | e.g.: line number 66-67: […] more detail in the Medical Devices Regulation (MDR) and In-Vitro Diagnostic Medical Devices Regulations (IVDR). -> […] more detail in the MDR and IVDR. | accept
intro | CA1 | 69 | ed, | We suggest to use the wording in accordance to Article 1 (1) IVDR. Therefore we propose to use the term Union market. | As a starting point, a medical device or an in-vitro diagnostic medical device may only be placed on the Union market if the specific requirements as defined in the Medical Device Regulation (EU) 2017/745 or in the In-vitro-Diagnostic Medical Device Regulation (EU) 2017/746 are met. | accept
1 | CA1 | 78 |  | We suggest to define the "non-scope" in addition to the "scope" in order to avoid a repeated formulation of what is not scope -> e.g.: guidance on trend reporting |  | reject
1 | CA1 | 93-94 | te |  | PMS for in vitro diagnostic medical devices according to IVDR is in principle equal to the PMS for medical devices according to MDR. | reject - to include legacy
1 | CA1 | 90 | ed | Typo in …(AIMDD) (90/385/EEC) or the In Vitro Medical Directive (IVDD) (98/97/EEC) according to | […] (AIMDD) (90/385/EEC) or the In Vitro Diagnostic Medical Devices Directive (IVDD) (98/97/EEC) according to […] | Removed
1 | CA1 | 100-101 | ed | typo | […] of medical device manufacturers. In particular, Article 10 (9) MDR and Article 10 (8) IVDR require manufacturers to establish, document, implement, maintain, keep up to date […] | accept
1 | CA1 | 104 | ed | “for” is missing | […] risk class and appropriate for the type of device. | for was added (per MDR)
2 | CA1 | 130-133 | te | Instead of redefining the term “medical device”, we propose to adopt the term "device" according to Article 1 (4) MDR and Article 1 (2) IVDR including relevant additional definitions in Article 2 IVDR, which are not part of Article 2 (2) (e.g.: device for self-testing, device for near-patient testing, companion diagnostic,…) |  | accept
2 | CA1 | 134-135 | ed | Delete sentence, as no added value |  | accept
3 | CA1 | 170 ff | te | Please define the difference of active and proactive approach for collection and review of information on post market experience (line 151 [pro-/active] vs 161 [active] vs. 170ff [proactive]) We suggest that it would be beneficial in section 3 of the guidance to define two sub-sections – one subsection that describes the (pro)active approach in further detail and one subsection that describes the reactive approach. This may also reduce repetition of information on p.5 and 6 of the guidance. Please also include respective IVDR references where the proactive approach is mentioned in the IVDR as well, to ensure better understanding. e.g. citation 63 of IVDR, Annex III Part A 1.b and Part B 4. and 5.1 IVDR to reference what is meant by “proactive” collection and evaluation of PMS data. |  | accept the need for revision
3 | CA1 | 193-196 | te | In our opinion it is important that the manufacturer describes in his internal procedures his strategic approach to ensure a comprehensive PMS. | We suggest to rephrase the sentence: “A comprehensive post-market surveillance system should be understood to include the strategy, clear structures, roles and responsibilities, procedures, processes and resources required to implement [...]” | accept the need for revision
3 | CA1 | 148 | te | Please include the word “recital” for better understanding: | 4 recital (80) MDR / (81) IVDR | recital 80 MDR and recital 81 IVDR
3 | CA1 | 201 | te |  | […] patient, user or other person. | accept
3 | CA1 | 203 | te | Please delete “s” | […] preventive and/or corrective action […] | accept
3 | CA1 | 158 | ed | typo | […] has been placed on the market. | accept
3 | CA1 | 165-168 | ed | We should focus on the PMS System (=Post Market Surveillance System) | The PMS system is intended to continuously monitor the results of the post-market phase to update pre-market activities according to Article 83 lit 3. a)-d) MDR / 78 lit 3. a)-d) IVDR and to identify and initiate appropriate measures according to Article 83 lit 3. e)-h) MDR / 78 lit 3. e)-h) IVDR when needed. | reject . Articles are already mentioned in previous part.
3 | CA1 | 170 | ed | typo | Annex XIV, Part B, Sections 5 and 6 MDR/Annex XIII, Part B, Sections 4 and 5 of the […] | accept
3 | CA1 | 182 | ed |  | A manufacturer shall not just wait for information to be sent to them (e.g. via customer complaints) but must […] | accept
PMS as an integral part o | CA1 | 228 |  | The PMS plan comes before the PMS System, therefore it should be mentioned at the beginning and not at the end |  | accept text revised
PMS as an integral part o | CA1 | 207-208 | ed | We suggest to use the same citation style among the document | The PMS System as an integral part of the manufacturer's quality management system referred to in Art. 10(9) MDR / Art. 10(8) IVDR | accept
PMS as an integral part o | CA1 | 211 | ed | The list of action items (bullet points) to be done by the manufacturer seems not to be complete. We miss some information on PMS reporting obligations (e.g. Art. 80 and 81 IVDR) and how the mfr. shall maintain effectiveness of the PMS process. If this list of bullet points is meant to be a non-exhaustive listing of information as an overview, please indicate that this list is meant as non-exhaustive to avoid confusion. | Either include : This means that a manufacturer: (non-exhaustive list) or also include information on reporting obligations and how to maintain effectiveness of PMS | accept
PMS as an integral part o | CA1 | 213 | ed | Please check the references in the whole document and outline them properly: E.g.: Line number: 208, 213, 264, 285, 400-405, 456, 467, 529 (heading in the table [Article XY applicable regulation), |  | accept
4 | CA1 | 234 | te | PMS System is part of the QM-System -> this should be represented in the graphic as well |  | agreed to either take the graphic out entirely or replace with a better graphic that takes this comment into account.
4a | CA1 | 263 | ed | typo | […] and be the linked with the […] | accept
4a | CA1 | 264 | ed | Reference to IVDR is missing | Annex I section 3 IVDR , Benefit-risk analysis accord. to section 1 and 8 of Annex I IVDR and reference to TD (Annex II 5. IVDR) | accept. Simplified.
4b | CA1 | 270 ff | te | There is a lack of description of further methods and instruments to fulfil the obligations of the MDR and IVDR. Integral components (e.g. PMS report) is missing. |  | accept, will be adressed in another section of the document.
4b | CA1 | 271 | te | We suggest to use the wording in accordance to Article 83. 3 (b) MDR, because the IFU is not design or manufacturing information. | b) Updates to the design and manufacturing information, including the instructions for use and the labelling […] | accept
4b | CA1 | 276 | te | For additional clarity. | […] data is considered as an input for design and further development of the medical device […] | accept
4c | CA1 | 292-299 | te | Since all points are taken from the MDR (only slightly changed and/or abridged) we would suggest that here the same order as in the MDR is used. |  | accept
4c | CA1 | 296 | te | We suggest to use here the exact wording from the referenced MDR and not make gratuitous changes. | ensuring the continued acceptability of identified risks the benefit-risk ratio and | accept
4c | CA1 | 297 | te | We suggest to use here the exact wording from the referenced MDR and not make gratuitous changes. | detecting emerging identifying and analysing emergent risks on the basis of factual evidence., | accept
4c | CA1 | 314 | te | We suggest also to add the last mentioned point in the referenced IVDR. | • identifying possible systematic misuse. | Reject, already in the text
4c | CA1 | 287 | ed | Add reference | Under the MDR Art. 2. (48) indent 4 clinically […] | Rejected. Guidance will provide references to denitions at the beginning of the guidance.
4c | CA1 | 311 | ed | typo | • identifying and analyzing […] | accept
4c | CA1 | 330-331 | ed | typo | […] be part of the clinical evaluation / performance evaluation […] | Reject - sentence has been restrcutured.
4d | CA1 | 351 | ed | typo | […] report. The post-market […] | accept
4e | CA1 | 356 | ed | We suggest to rephrase this paragraph. [remark: repetition of information from line 207 to 231] | Manufacturers should set up their PMS system to actively collect and thereby analyse and utilize the relevant data on quality, performance and safety of the device throughout the entire lifetime. By this monitoring the manufacturer shall draw conclusions whether implementation of preventive and corrective actions are indicated. Furthermore the manufacturer shall set up processes and procedures to integrate the data obtained by the PMS system into the QMS, in order to initiate and document preventive and corrective actions as well as field safety corrective actions (FSCAs) to reduce the risk posed by products already placed on the market. | accept
4f | CA1 | 366-385 |  | Please change the order of f) and g) for better readability | f) Contribute to the post-market surveillance of other devices – Art. 83.3 (g) MDR Art. 78. 3 (g) IVDR g) Identify options to improve usability, performance, and safety of the device – Art. 83.3 (f) MDR Art. 78. 3 (f) IVDR | reject - this chapetr follows the order of art 83.3 and art 78.3
4f | CA1 | 372-374 | ed | typo | […] if there are any options to improve the safety or performance of the device and the feedback could be used to improve the use ease of the medical device. | partially accepted
4g | CA1 | 378-385 | te | We suggest to rephrase this section because it is difficult to understand |  | accept paragraphs are rephrased
5 | CA1 | 399-420 |  | The Plan comes for the established System (PMS System), therefore we would suggest to change the order for better understandability. |  | reject - the plan is part of the PMS
5 | CA1 | 400 | ed | Among the whole document there is no figure with the figure description “figure 1” Please add correct figure descriptions. |  | accept
5.1 | CA1 | 408 | Ed | Please delete “for” | “…should be based on the requirements for which are set out in Section 1 …” | accept
5.1.1 | CA1 | 414 | ge | The line "Information, including feedback and complaints is [...]" is included twice in the table. We suggest to leave only the line on page 13 and change the text as indicated to the right. | The post-market surveillance plan should explain how information provided by users, distributors and importers feedback and complaints are is received […] Manufacturers are expected to be proactive in gathering information, feedback and complaints. […] | acepted
5.1.1 | CA1 | 414 | Te | Please reference to the corresponding documents in a uniform way |  | accept
5.1.1 | CA1 | 414 | Ed | Please delete the second “that” | “…the methods and tools that that are used…” | accept
5.1.2 | CA1 | 419 | Te | Please correct the name of the referenced Q&A on trend reporting and align the foot note in the table with the number at the foot line of the page: The post-market surveillance plan should describe the manufacturer’s processes and methods for investigating complaints and how to collect and analyse data related to market-related experience collected in the field.. More about trend obligations you can consult the MDCG Q&A trend 8 .refer to the trend Q&A |  | tabel is revised
5.2 | CA1 | 422-485 |  | More or less repetition of the content already mentioned above |  | this is the
5.2 | CA1 | 433-436 | ge | We don’t know exactly what this refers to. What is the manufacturer required to do? Are usability data meant? |  | accept
5.2 | CA1 | 422-485 |  | More or less repetition of the content already mentioned above |  | accept
figure 437 | CA1 | 438 | Ed | Figure 2 is a very helpful overview: we suggest to add in the bubble “user feedback as an additional example for surveys to include also “user surveys” as this survey are different from customer surveys (as the customer is not always the user). Those surveys can e.g. either be written or online In the bubble Post Market Clinical Studies: please correct PMCF/PMF to PMCF/PMPF studies. | PMCF/PMPF studies | accept
figure 437 | CA1 | 438 | Ed | Figure 2 is a very helpful overview: we suggest to add in the bubble “user feedback as an additional example for surveys to include also “user surveys” as this survey are different from customer surveys (as the customer is not always the user). Those surveys can e.g. either be written or online In the bubble Post Market Clinical Studies: please correct PMCF/PMF to PMCF/PMPF studies. | PMCF/PMPF studies | accept
figure 437 | CA1 | 438 | Ed | Figure 2 is a very helpful overview: we suggest to add in the bubble “user feedback as an additional example for surveys to include also “user surveys” as this survey are different from customer surveys (as the customer is not always the user). Those surveys can e.g. either be written or online In the bubble Post Market Clinical Studies: please correct PMCF/PMF to PMCF/PMPF studies. | PMCF/PMPF studies | accept
5.3 | CA1 | 462 | ed |  | “…actions (ANNEX III 1 (b) MDR or IVDR). | accept
5.3 | CA1 | 462 | ed |  | “…actions (ANNEX III 1 (b) MDR or IVDR). | accept
5.4 | CA1 | 479 | te | Article 85 refers to PMS report and article 86 refers to PSUR. An additional "summary report" relating to PMS is not defined in the PMS chapter. It could also lead to confusion with the "perodic summary report" described in article 87 (9). | […] be drawn up in a summary report. the PMS report or the PSUR89. | accept
5.4 | CA1 | 476 | ed | typo | So besides planning the PMS, and follow the procedures for collecting information. The the data needs to be analysed and assessed. […] | accept
5.4 | CA1 | 482 | ed | typo | to To start the next period […] | accept
5.4 | CA1 | 482-485 | ed | We suggest to rephrase the paragraph to ensure proper interpretation. | In order to start the next period of PMS a review of the existing system is necessary. Within this review it is determined if the plan is still up to date, if proper data sources are still in use, if the collected information addresses the objectives of the PMS and if the findings are used within the relevant processes. | accept
5.4 | CA1 | 482-485 | ed | We suggest to rephrase the paragraph to ensure proper interpretation. | In order to start the next period of PMS a review of the existing system is necessary. Within this review it is determined if the plan is still up to date, if proper data sources are still in use, if the collected information addresses the objectives of the PMS and if the findings are used within the relevant processes. | perhaps make clear that PMS system is a plan do check act? Then 5.2, 5.3 and 5.4 are mor logic?
6 | CA1 | 500 | ge | We suggest to mention here the obligation of a PMS report for class I custom-made devices. | Manufacturers of class I custom-made devices shall prepare a PMS report summarising the results and conclusions of the analyses of the post-market surveillance data gathered as a result of the post-market surveillance plan. | accept
6 | CA1 | 496 | ed | typo | […] it learns of them.’ | accept
6 | CA1 | 496 | ed | typo | […] it learns of them.’ | accept
6 | CA1 | 496 | ed | typo | […] it learns of them.’ | accept
annex I | CA1 | 528-530 | Ed | We suggest to delete “articles of the regulations” Please use the same language style (british/US) among the document Please add “ratio” | “PMS System – MDR Article 83 / IVDR Article 78Articles of the Regulations” Analysing “…reassessing the benefit risk ratio.” | accept
1 | CA2 | 91-97 | TE | “Requirements for IVD Legacy devices have been detailed in MDCG 2022-8.” There exists a similar guidance for MDR legacy devices namely MDCG 2021-25 | Add a reference to MDCG 2021-25 | accept propose to only add wording for the legacy by referring to the MDCGs
2 | CA2 | 126-135 | GE | It does not make sense to redefine the term medical device. The regulations use for this the term device see article 1 paragraph 4 of the MDR and article 1 paragraph 2 of the IVDR | Use the term device instead of medical device | Accept - it is not the definition but how to read 'medical device' in this guidance
3 | CA2 | 145-148 | GE | “National competent authorities within the European Union are obliged to ensure effective coordination of their market surveillance activities and to clarify the applicable procedures.” It is better to delete the “withing the European Union” as also national competent authorities of Norway, Iceland, Liechtenstein & Turkey need to do this. | “National competent authorities are obliged to ensure effective coordination of their market surveillance activities and to clarify the applicable procedures.” | part regarding CA will be deleted not for guidance
4c | CA2 | 311 | ED | “• dentifying and analysing emergent risks on the basis of factual evidence” Missing an I | • Identifying and analysing emergent risks on the basis of factual evidence | accept
4c | CA2 | 330 | ED | “.. clinical evaluation r/ performance evaluation report and the technical documentation.” the r before r/ | “ clinical evaluation / performance evaluation report and the technical documentation.” | accept
4d | CA2 | 349-352 | ED | “Such updates should be aligned with the information presented in the clinical evaluation report and Periodic Summary Update report The post-market surveillance system should describe how the PMS data is used to update the SS(C)P, when applicable.” It is Periodic Safety Update Report. A point is also missing | “Such updates should be aligned with the information presented in the clinical evaluation report and Periodic Safety Update Report. The post-market surveillance system should describe how the PMS data is used to update the SS(C)P, when applicable. | accept
5.1.1 | CA2 | 414 | TE | In the “ How to collect it” column in row “ Publicly available information about similar medical devices” We would include EUDAMED in this list, as there is publicly available information from other similar devices like the SS(C)P and vigilance data | - Add : EUDAMED | accept
5.1.1 | CA2 | 414 | TE | In the “ How to utilize it” column in row “ Publicly available information about similar medical devices” “The data that is gathered via the activities listed within Annex III 1.1(a) should be used to monitor any emerging risks as well as to help demonstrate continued acceptability of the performance and safety of the device.” There is no 1.1(a) in that Annex; it should be Annex III 1. (a). see latest version of the MDR: https://eur-lex.europa.eu/legal-content/EN/TXT/HTML/?uri=CELEX:02017R0745-20230311&from=EN | “The data that is gathered via the activities listed within Annex III 1. (a) should be used to monitor any emerging risks as well as to help demonstrate continued acceptability of the performance and safety of the device.” | accept
5.1.1 | CA2 | 414 | ED | In the “ How to collect it” column in row “Information, including feedback and complaints, provided by users, distributors and importers;” “The post-market surveillance plan should address how feedback and complaints are received and processed by the manufacturer, as well as the methods and tools that that are used to investigate complaints and to analyse market-related experience collected in the field. .” Point too many at the end of the sentence | “The post-market surveillance plan should address how feedback and complaints are received and processed by the manufacturer, as well as the methods and tools that that are used to investigate complaints and to analyse market-related experience collected in the field.” | accept
5.1.1 | CA2 | 414 | ED | In the “ How to collect it” column in row “Information from trend reporting” “This includes those that could have a significant impact on the benefit-risk analysis in relation to the GSPRs and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits.” GSPR is not explained previously. We would simply write requirements of annex I | This includes those that could have a significant impact on the benefit-risk analysis in relation to the requirements of Annex I and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. | accept
5.1.2 | CA2 | 419 | ED | “— reference to procedures to fulfil the manufacturers obligations laid down in Articles83, 84 and 86 MDR and article 78, 79 and 81 IVDR;” “The manufacturer should clearly identify the relevant procedures for post-market surveillance activities as described in Article 78, 79 and 81, including the PMS plan, the data to be collected and recorded, its analysis and the requirements for the PMS report / PSUR. The procedures for actions to be taken including corrective actions should be identified.” In the first column there is an editorial error as there is a space missing before 83. The issue is mainly that in the whole document the MDR references are used or both while the second column only mentions the IVDR references. | “— reference to procedures to fulfil the manufacturers obligations laid down in Articles 83, 84 and 86 MDR and article 78, 79 and 81 IVDR;” “The manufacturer should clearly identify the relevant procedures for post-market surveillance activities as described in Article 83, 84 and 86 MDR and Article 78, 79 and 81 IVDR, including the PMS plan, the data to be collected and recorded, its analysis and the requirements for the PMS report / PSUR. The procedures for actions to be taken including corrective actions should be identified.” | accept
5.3 | CA2 | 462 | TE | “The post-market surveillance plan should detail or reference the procedure in relation to these corrective actions (ANNEX IIIb).” The reference to Annex III is not correct and not consistent with previous usage | The post-market surveillance plan should detail or reference the procedure in relation to these corrective actions (Annex III 1. b). | accept
5.4 | CA2 | 480 | ED | “Conclusions need to be drawn up in a summary report. the PMS report or the PSUR8.” Not sure if the 8 is a typo or that a footnote is missing. In case it is a typo remove the 8. | “Conclusions need to be drawn up in a summary report. the PMS report or the PSUR.” | accept
5.4 | CA2 | 482-486 | ED | “to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes.” Something went wrong with the interpunction in these lines | To start the next period of PMS, a review of the PMS plan is necessary. The purpose of this review is to determine if the plan is still up to date, if the proper data sources are still in use, if the collected information address the objectives of the PMS plan and whether the findings are used for the relevant processes. | accept
annex I | CA2 | 530 | ED | Last row of table: “Implantable and Class III devices: • Updated when necessary and at least annually. • Submitted by means of EUDAMED to notify body. • NB evaluate and add its evaluation report in EUDAMED.” Missing “• PSUR and NB evaluation available to CA in EUDAMED” as also these are available to the CA not only the IVDR class D PSUR | “Implantable and Class III devices: • Updated when necessary and at least annually. • Submitted by means of EUDAMED to notify body. • NB evaluate and add its evaluation report in EUDAMED.” “• PSUR and NB evaluation available to CA in EUDAMED” | Text was revised.
3 | COCIR | 137 | ge | It is necessary to also describe the need of creating product families and how they have to align with technical document. At minimum a reference to the MDCG guidance describing how to create product families should be made |  | reject - the question of grouping is valid but the guidance is not the place to discuss; maybe a Q&A will be required
grafic 1 | COCIR | 241 | ge | Missing a reference to “State of the Art” |  | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | COCIR | 262-264 | ge | “Identification of thresholds for trend reporting is not always feasible/possible. specially for medical devices falling in the lower risk classes. In particular when the compliant rate, registered in the observational period, is limited in quantity and heterogenous in nature, statistical-based conclusion on “significant increase or severity” of the events , as per MDR art 88, nor trend reporting can be conducted. In all these cases individual root-cause analysis should be preferred. | “Whenever feasible, indicators and threshold values[1] used in the continuous reassessment of the benefit-risk analysis of the device should be covered in the PMS Plan or a proper justification for no trend reporting should be provided by the manufacturer Such trending data should be the linked with the risk management documentation (as referred to in Section 3 of Annex I MDR | rejected - feasible is not aligned with the MDR. Instead, we will introduce the wording "suitable".
4a | COCIR | 262 | ed | Footnote: “guidance Q&A on trend report” is unclear and does not help the reader to find the referred document | Insert a clear reference to MDCG document or remove the footnote | accept
4b | COCIR | 273 | ge | IFU and labeling are required per MDR, but would it make sense to add into this category user training? Based on the outcome of the data analysis it might be necessary to improve the user training to avoid risks and/or misuse | Information obtained through the post-market surveillance process should be used to update the design and manufacturing information, user training information (if applicable), the instructions for use and the labelling. | accept. Extend the list of examples: surgical techniques, user training, user training as part of installation training.
4c | COCIR | 321 | ge | The clinical evaluation should be described in more details, especially that there is a general approach to clinical evaluation which is required always which entails literature review, review of registries and other data bases, feedback from users etc. the need for specific clinical evaluation has to be assessed and if not required explained why not required. | General methods and procedures of PMCF/PMPF should include post-market activities, such as gathering of clinical experience gained, feedback from users, screening of scientific literature, evaluation of suitable registers. Specific methods include PMCF/PMPF studies. | Reject. TF is of the opinion to stay away from repeating what has already been described in other MDCG guidance. Therefore, document will refer to already establied MDCG guidance. This guidance will not explain in detail how to update the clinical evaluation.
4c | COCIR | 311 | ed | Spelling, currently “dentifying” should be “identifying” | identifying | accept
4c | COCIR | 330 | ed | Difficult to understand and possible spelling error? “The manufacturer shall analyse the findings of the PMCF/PMPF and document the results in a PMCF/ PMPF evaluation report that shall be part of the clinical evaluation r/ performance evaluation report and the technical documentation” |  | Reject.
5 | COCIR | 403 | ed | Typos: “In paragraph 6.1.1” | In paragraph 5.1.1” | accept
5 | COCIR | 405 | ed | Typos: “In paragraph 6.1.2” | In paragraph 5.1.2” | accept
5.1 | COCIR | 411 | ge | “publicly available information about similar devices » Should this not include the reference to “equivalent“ devices? | publicly available information about similar and equivalent devices | reject
5.1 | COCIR | 411 | ed | Line item 1 and line item 6 are duplicated, the descriptions should be merged |  | accept
5.1.2 | COCIR | 419 | ge | Annex III indicates justification could be provided as to why PMCF may not be applicable, however, it’s not apparently clear when PMCF may be not applicable. (specifically general activities). | Provide examples within the guidance doc of acceptable circumstance(s) in which a manufacturer could determine PMCF activities are not required. | reject. This should be clarified in a MDCG Guidance for PMCF.
5.1.2 | COCIR | 419 | ge | Suitable indicators and threshold values: Under explanation it would be necessary to indicate that the original set of indicators and thresholds have to be defined by the premarket phase. The threshold and indicators can then be updated as part of the review and the collection and analysis of the post market data | Suitable indicators and thresholds shall be defined within the pre market stage and should be used for pre market verification and validation activities as required. During the post market phase the thresholds and indicators require to be reassessed for their validity. The data obtained from these activities should also be used to set thresholds and indicators that will allow for the continuous reassessment of the safety and performance of the device and also enable trend reporting. | tabel is revised
5.2 | COCIR | 482 | ge | Add reference to triggers and thresholds | To start the next period of PMS, a review of the PMS plan is necessary. This review is to determine if the plan is still up to date, if the proper data sources are still in use, if the identified triggers and thresholds are still appropriate, if the collected information addresses the objectives of the PMS plan, and whether the findings are used for the relevant processes | accept
figure 437 | COCIR | 436 | ge | What about manufacturing and reliability data? |  | reject
4c | COCIR | 30 and 285 | ed | Wrong reference to MDR Art 83 “Art. 83. 3 (b) MDR” Similar problem on IVDR | “Art. 83. 3 (c) MDR” | accept
ge | CA3 | Whole document | ge | General comment: Due to its structure, the draft contains many redundancies leading to an unnecessary length. What is missing in the document is a clear description of the interrelation of the various processes, i.e. which processes should be seen as (sub-) processes of more general processes like risk-management or the clinical evaluation. This is (partially) mentioned in section 1 Scope and Objective (lines 82-83, 98-125) but in the text, this objective isn’t appropriately met. Note: MDR Art. 19 section 9 does not provide clear requirements on how which of the required processes are related but this guidance could be a good place to provide such clarification, e.g. by outlining a process landscape at least for the (sub-)processes covered, e.g. how exactly PMS is linked to risk management and clinical evaluation. In this respect, a clear input / output approach (see comment below) could also be helpful. | Suggestion to streamline the content in a way that the manufacturer’s obligations will be explained once with a better description of the interrelation of the various processes, i.e. to combine chapters 3 and 4 and Annex I as well as including into this new chapter parts of chapter 5. Aim should be that the tasks of a manufacturer are described once and comprehensively within this description. | accept the need for revision
ge | CA3 | 137-508 | ge | See general comment In principle, the approach taken in chapter 4 (figure on page 7) is understandable but it would be more helpful when the interrelation of processes would be displayed e.g. as process landscape to better describe their interrelation. E.g. in lines 243-269, a clear link / relation to the risk-management process is missing. Also a clear and consistent indication of inputs – outputs could be helpful. This is partially – but not consistently – applied and as a description of the “logic” behind the structure (chapters 3-5, Annex I) is not provided the guide triggers unfortunately many questions. Chapter 5 contains a kind of such a “serial approach” but as explained above it would be more helpful to apply only one (comprehensive) description, at least to combine the graphics on page 7 and 16 with a clear indication of inputs / outputs. When the graphic on page 7 will be kept, at least reference to the respective sections (letters a) to h) should be incorporated. In addition, references to other guidance documents (e.g. SCCP) should be made. | For discussion | accept the need for revision
ge | CA3 | Whole document | ed | Editorial aspects: The draft currently contains a series of inconsistencies (e.g. graphics / figures are not numbered consistently, abbreviations are not introduced once and used consistently later on, footnote numbers are not correct), typos etc. which should be corrected in the next version. | Please check and correct accordingly | accept
1 | CA3 | 91 | ed | the amending regulation (EU) 2023/607 should be mentioned | in connection with Reguation (EU) 2023/607 | ask Hanim how to refer to this - should it be incorporated?
2 | CA3 | 126-135 | te | What is the reason for defining “medical device” when already the Regulations deal with this, see e.g. MDR Art. 1(4) and IVDR Art. 1(2). Making reference to definitions in “existing MDCG guidance’s” should be avoided as in some guidance “deviating definitions” being applicable only for the purpose of such guidance are used. The last paragraph should be deleted: When a definition is needed then it should be described in chapter 2. Otherwise such a chapter wouldn’t be needed. | Suggest to refer to those Articles and keep the same terminology. Please change or delete the second sentence (lines 128-129) as well as the last paragraph (lines 134-135). | Accept - it is not the definition but how to read 'medical device' in this guidance
3 | CA3 | 165 | ed | Monitoring of the pre-market phase ends by placing the product on the market. Therefore “verify” would be more appropriate | Replace “monitor” by “verify” | accept
grafic 1 | CA3 | 242 | ge | Figure shows activities related to PMS. However, since most of the listed aspects are dependent on each other, proposal is to visualize these dependencies to get a clear picture of the process/information flow from PMS as input to the deducted measures | Aspects being related to PMS are simply listed. Proposal to adapt figure to visualize dependencies in a process flow chart. E.g. PMS delivers data to update Risk assessment + clinical evaluation + trendingà Risk assessment, trending and clinical evaluation deliver data to update benefit-risk determination and to identify measures/CAPA such as update/improve of IFU/label/design/manufacturing information/usability/performance/safety | agreed, graphic will be taken out or replaced with a different graphical representation.
grafic 1 | CA3 | 242 | te |  | Include footnotes to explain abbreviations used in figure | agreed, graphic will be taken out or replaced with a different graphical representation.
grafic 1 | CA3 | 242 | ed | I understand that the central goal of PMS is to periodically update the “new” or “reassessed” risks in the risk assessment rather than improving the risk management process itself. Therefore “update” would be the more appropriate term than “improve” and the order would be 1) update risk assessment 2) update benefit-risk determination | Upper oval shape in the figure: Replace “update benefit-risk determination & improve risk management” by “update risk management & benefit-risk determination” | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | CA3 | 243-249 | ed | I understand that the central goal of PMS is to periodically update the “new” or “reassessed” risks in the risk assessment rather than improving the risk management process itself. Therefore “update” would be the more appropriate term than “improve” and the order would be 1) update risk assessment 2) update benefit-risk determination | Change sequence of the described activities: 1) update risk assessment 2) update benefit-risk determination”. Change wording “improvement of risk management documentation” to “update of risk management documentation” | accept, section will be re-written and this comment will be considered.
4c | CA3 | 296 | te | Advice should be added according to Annex XIV, Pt. B, 6.1d: This is the basis to monitor also the “state of the art” treatments for the indication / intended purpose (according to Annex I, chapter I, 1), which determines the current benefit of the product. | Add “and benefit-risk ratio” to the bullet point: “ensuring the continued acceptability of identified risks and benefit-risk ratio and” | accept
4c | CA3 | 324 | te | Supplement sentence to give guidance to monitor also the “state of the art” treatments for the indication / intended purpose (according to Annex I, chapter I, 1) | Add to the sentence “and collecting current information regarding alternative treatment options/products on the market for the same intended purpose.” | Reject, the relevant part has been removed from the guidance.
4c | CA3 | 332-333 | te | The sentence says that “PMCF/PMPF evaluation report shall be taken into account for the clinical evaluation/performance evaluation”. This is misleading since the sentence before says that “results in a PMCF/ PMPF evaluation report that shall be part of the clinical evaluation r/ performance evaluation report “ | Delete “for the clinical evaluation/performance evaluation and” | accepted, this part of the guidance will be rephrased.
4c | CA3 | 287-290 | ed | Sentence is unclear (syntax) | Rephrase | accept, has been rephrased.
4f | CA3 | 372 | ed |  | Add “are” in “…if there “are” any…” | accept
4g | CA3 | 384 | ed | Impact on other devices shall be assessed rather than the impact on other users | Change wording from “limit the impact to other users of the device” to “limit the impact to users of the other device | accept paragraphs are rephrased
5 | CA3 | 399-485 | te | See general comment above In the tables on pages 11-15 some references / indication where the text is coming from, is missing, e.g. the second line in table 1 “Information, including feedback and complaints, …” | We suggest to combine the various sections where manufacturer’s tasks are described to only one section with a consistent and “explained logic” which should be described in an introductory part of this section. What needs to be avoided are expressions like “There is an expectation that manufacturers have …”. Either there is a requirement or not. If not, then we should not “expect” something. | comment will be used when reviewing the whole table
5.1.1 | CA3 | 414 | te | Line 3, column 3 Delete utilization for trend reporting: Trend reporting is supposed for non-serious incidents and this line addresses serious incidents | Delete “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR.“ | accept
5.1.1 | CA3 | 414 | te | Line 4, column 3 Paragraph is saying that trending data is used to identify new risks. But since trending is performed on known risks, there should not result new risks, but the assessment of the known risks could change | Change paragraph: Trending data is used to reassess whether known risks are still non-serious and acceptable. | accept
5.1.1 | CA3 | 414 | te | Line 5, column 3 Paragraph is saying that trending data is used to identify new risks and it should be considered to change the labelling accordingly. But since trending is performed on known risks, there should not result new risks, but the assessment of the known risks could change | Delete paragraph: “The manufacturer should also consider the need to update the labelling and where applicable update to the SS(C)P to inform users of the device of any new or emerging risks.” | accept. See also comment in line 473.
5.1.1 | CA3 | 414 | te | Line 8, column 3 Add clause to provide guidance on what would be the goal of comparison with other products. | Complement 2nd sentence with “and thereby to continuously assess the benefit risk ratio.” à “Annex III also highlights that this data should be used to enable comparison of device with similar products on the market and thereby to continuously assess the benefit risk ratio.” | rejected.
7 | CA3 | 513-514 | te | Standards should be cited in the correct version, especially when there is a harmonized standard. We also wonder why only “clause 8” is stated here when this standard relates also to the manufacturer’s QMS in general. | EN ISO 13485:2016/A11:2021 | partially accept. For the title of the ISO standard, see COMMISSION IMPLEMENTING DECISION (EU) 2022/757
intro | CA4 | 66-67, etc, | ge | In line 66-67, the following is written; Medical Devices Regulation (MDR) and In-Vitro Diagnostic Medical Devices Regulations (IVDR). As IVDR and MDR are listed as abbreviations in line 58-59, please continue to use these abbreviations throughout the entire document to ensure consistency. | Use MDR and IVDR | accept
intro | CA4 | 69 | ed | Please consider using “when” instead of “if”, as “if” indicates that this must be fulfilled once, whereasaccept using “when” indicates that this must be continually fulfilled and that it must be continuously evaluated whether criteria are met or not. | Replace “if” with when” | accept
1 | CA4 | 100-104 | te | Please include that the QMS contains more than the PMS system, to emphasize that the QMS contains more obligations than the PMS system. | Add text about QMS | accept
1 | CA4 | 102 | ed | Please include “QMS” as abbreviation for quality management system | Add (QMS) after quality management system | accept
2 | CA4 | 134-135 | ge | Please remember to update with footnote or reference. | Add footnote/reference | accept
3 | CA4 | 153 | te | Consider adding a reference to MDR article 83 and IVDR article 78 | Add reference | already in scope and objective
3 | CA4 | 138 | ed | Rephrase ’Surveillance start at the moment that activities are initiated when the’ | Rephrase to ‘Surveillance activities are initiated when the’ | accept
3 | CA4 | 141 | ed | Delete “” and “” in the sentence “between “market surveillance” activities and “post-market surveillance” activities”. Possibly cursive. | Rewrite | accept
3 | CA4 | 167 | ed | Add “and for the manufacturer” to the sentence | Rephrase to; “…experience gained during the post-market phase and for the manufacturer to identify…” | accept
3 | CA4 | 172 | ed | Please elaborate on PMPF as this abbreviation appears here for the first time in the document | Add Post market performance follow-up before PMPF | accept
3 | CA4 | 204 | ed | “the” is missing in the sentence | New wording “relevant part of the technical documentation” | accept
PMS as an integral part o | CA4 | 208 | ed | Add “MDR” | Article 10(9) (IVDR 10(8)) à MDR article 10(9) (IVDR 10(8)) | accept
PMS as an integral part o | CA4 | 214 | ed | Instead of referring to the “bullet point above” – please elaborate where (section) for clarification | Refer to section/subparagraph/page or similar | accept
grafic 1 | CA4 | 242 | ed | Consider adding the adjacent letter sectioning to the corresponding bobbles/circles of the figures to enhance the comprehension | Add a), b), c) etc. to figure/bobbles/circles. | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | CA4 | 253-254 | ed | New wording of sentence to correct grammar; • the planning, conduct, assessment and reporting of post-market surveillance activities | • the planning, conducting, assessing and reporting of post-market surveillance activities | accepted
4a | CA4 | 263 | ed | Delete “the” in “and be the linked to the risk” | à “and be linked to the risk” | accept
4b | CA4 | 279 | Ed | ‘s’ is missing | Contradiction à contradictions | accept
4b | CA4 | 282 | Ed | Delete “to” in “…or to changes to the design” | à “…or changes to the design” | accept
4c | CA4 | 284-338 |  | Suggestion: add a figure showing how the PMCF/PMPF and PMS are intertwined, and how the input/output of these feeds into each and further onto risk management (eg. evaluation of benefit-risk ratio). | Add figure | will be considered as part of re-draft of other graphics as mentioned in earlier comments above.
4c | CA4 | 302 | te | Should it be “post market performance follow-up” instead of “post-market clinical follow-up” |  | accept
4c | CA4 | 311 | ed | I is missing | Dentifying à Identifying | accept
4c | CA4 | 311 | ed | Use analysing or analyzing in a consistent matter throughout the entire document | Choose one of two spelling options | accept
4c | CA4 | 330 | ed | Delete “r” in “ evaluation r/” | Delete r in “evaluation r/” | Reject - sentence has been restrcutured.
5 | CA4 | 403 | Ed | Change reference 6.1.1 to 5.1.1 | Change reference 6.1.1 to 5.1.1 | accept
5 | CA4 | 405 | ed | Change reference 6.1.2 to 5.1.2 | Change reference 6.1.2 to 5.1.2 | accept
5.1.1 | CA4 | 414 | te | Consider adding “e.g. relevant international and clinical databases” after the sentence “The manufacturer is expected to screen the relevant scientific, technical, and clinical literature and registers” |  | accept, text adjusted.
5.1.1 | CA4 | 414 (table) | te | Since FSN’s are required as part of an FSCA, delete “which could include the publication of a field safety notice” in “When trends are identified or there is evidence of a new or emerging risk, consideration should be given to whether field safety corrective actions are required, which could include the publication of a field safety notice.” | à When trends are identified or there is evidence of a new or emerging risk, consideration should be given to whether field safety corrective actions are required. | accept - deleted
5.1.1 | CA4 | 414 (table) | te | Consider to add “continuous trending on complaints with set trigger points” after ,”field” in “ The post-market surveillance plan should address how feedback and complaints are received and processed by the manufacturer, as well as the methods and tools that that are used to investigate complaints and to analyse market-related experience collected in the field.” |  | reject - "trigger points" is not in line in the Regulations.
5.1.1 | CA4 | 414 | ed | “Information, including feedbacks and complaints, provided by users, distributors and importers” is listed twice under “Required Information as per Annex III 1.a” |  | accept
5.1.2 | CA4 | 419 | ed/te | Please make sure that the MDCG Q&A trend is published before referring to it. |  | Noted.
figure 437 | CA4 | 437 (figure 2) | te | In figure 2, complaints are placed together with serious incidents. In the above table (5.1.1.) the complaints are placed together with user feedback. | Suggestion to choose only one grouping and use that throughout the guidance document. | accept, figure is deleted
5.4 | CA4 | 482-485 | ed | Consider rewording of “to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes.” And move to the beginning of the section | New wording; Prior to initiating the next data collection period of PMS, a review of the current PMS plan is necessary. This review is to determine if the plan is still up to date, if the proper data sources are still in use, whether the collected information address the objectives of the PMS plan, whether the findings are used for the relevant processes, or if changes to the PMS plan are needed. | accept
6 | CA4 | 494 | ed | change “it” to “the manufacturer” | à “In that context, the manufacturer shall report” | accept
6 | CA4 | 496 | ed | Change “..as soon as it learns of them.” To “..as soon as they are made aware” | Change “..as soon as it learns of them.” To “..as soon as they are made aware” | accept
7 | CA4 | 512 | ed | Remember to choose wording | Choose one og two wordings | accept
annex I | CA4 | 529 (table) | ed | Under table section “PSUR - MDR Article 86 / IVDR Article 81” notify body is used twice | Change “notify body” to “notified body” | accept
4d | CA4 | 351 | te | Consider adding ”…as well as the Instructions for Use and labelling” after “report” | à …Periodic Summary Update report, as well as the Instructions for Use and labelling. | Reject - text was taken out.
3 | CA4 | 147 | te | Consider adding reference to MDR article 93-100 and IVDR article 88-95 | Add reference | part regarding CA will be deleted not for guidance
PMS as an integral part o | CA4 | 224 | te | Add ”preventive and field safety corrective actions” to the sentence | New sentence; “… appropriate measures including corrective, preventive and field safety corrective actions if effects are identified that could…” | accept text revised
1 | CA4 | 108 | ed | Please add “post market clinical follow up” before PMCF for this abbreviation as it appears for the first time in the document | “including PMCF” à “including post market clinical follow up (PMCF)” | accept
ge | EAAR |  | Ge | · There is an overlap between the requirements for PMS and PMPF: a clear guidance what is understood to be documented in PMPF and what in PMS is lacking. E.g. Annex III, 1(a) states publicly available information about similar devices relevant specialist or technical literature, databases and/or registers Annex XIII, 5.2 lists (among others: Screening of scientific literature, Evaluation of suitable patient or disease registers, genetic databanks Evaluation of performance data relating to equivalent of similar devices | In our opinion, PMPF builds the bridge between the clinical evidence collected pre-market and the performance when the device is in routine use. It is the intersection between Performance Evaluation and Post Market Surveillance. Therefore, data relating to the clinical evidence of the device (for IVDs: scientific validity, analytical performance and clinical performance), needs to be planned and reported as part of PMPF. This can include literature studies, database registries, ring trial reports, PMPF studies, …
ge | EAAR |  | Ge | We feel that the concept and importance of State of the Art is not sufficiently emphasized: it is crucial that during the lifecycle of a medical device, it remains state of the art in terms of performance and in terms of current clinical practice. |  | accept
intro | EAAR | 64ff | ed | There are a few double-spaces in the document. | Search for and correct double-spaces. | accept
intro | EAAR | 66ff | ed | There may be too many introductions of the abbreviations MDR and IVDR. One is in line 66 and another one in line 98. In between (line 70) is the use of the full term, although the abbreviation was already introduced. There are also different wordings like “Medical Device Regulation” or “Medical Devices Regulation”. The terms “In-Vitro”, “in vitro” and “In Vitro” are used. In summary, page 3 is a bit inconsistent. | Revise inconsistencies. | accept
1 | EAAR | 88 | te | The term “certified” may be misleading, because there are legacy devices without Notified Body certificate. PMS requirements also apply to those. | Use different term than “certified under”. Maybe “covered by”. | accept
1 | EAAR | 104 | ed | Change “appropriate the type” to “appropriate to the type”. | Change “appropriate the type” to “appropriate to the type”. | reject (see above)
3 | EAAR | 139/140 | ed | Change “specified in the regulation” to “specified in the applicable regulation”. | Change “specified in the regulation” to “specified in the applicable regulation”. | accept
3 | EAAR | Footnote 4 | ed | It may not be clear that the footnote is referring to the recitals. The IVDR reference is missing. | Change to “Recital (80) MDR / Recital (81) IVDR”. | accept
PMS as an integral part o | EAAR | 213 | ed | Remove “. (“ at the end of the line. | Remove “. (“ at the end of the line. | accept
4a | EAAR | 248 | ed | Remove duplicate “the”. | Remove duplicate “the”. | accept
4c | EAAR | 329-330 | Ge | Document states “… in a PMCF/ PMPF evaluation report that shall be part of the clinical evaluation r/ performance evaluation report”: Since PMPF is part of PMS, it is warranted that the PMPF report is part of the PMS report/PSUR (as is the PMPF plan part of the PMS plan as per Annex III (b) last indent). The information/conclusions in the PMPF report shall then be used to update the performance evaluation report as per Annex XIII, 6. | “… in a PMCF/ PMPF evaluation report that shall be part of the PMS report/PSUR”: | Reject, this part of teh gudiance in not the right place to point to the PMSR/PSUR.
4c | EAAR | 288 | ed | Change “considered “clinical data”, Annex XIV” to “considered “clinical data”. Annex XIV” (the comma should be a full-stop). | Change “considered “clinical data”, Annex XIV” to “considered “clinical data”. Annex XIV” (the comma should be a full-stop). | accept, has been rephrased.
4c | EAAR | 330 | ed | Change “r/” to “report /” at the end of the line. | Change “r/” to “report /” at the end of the line. | Reject - sentence has been restrcutured.
4d | EAAR | 351 | ed | Add full-stop directly after the first word “report”. | Add full-stop directly after the first word “report”. | accept
4f | EAAR | 372 | ed | Change “if there any options” to “if there are any options”. | Change “if there any options” to “if there are any options”. | accept
5 | EAAR | 403 + 405 | ed | The correct paragraph numbers seem to be “5.1.1” and “5.1.2”. | The correct paragraph numbers seem to be “5.1.1” and “5.1.2”. | accept
5.1.2 | EAAR | Section 5.1.1 and 5.1.2 | Ge | The “how to use it” and “explanations” in Section 5.1.1 and 5.1.2 are theoretical. Examples would improve the guidance document. |  | Noted - It is the intention of the TF to provide examples in a future update but not as part of the initial release.
5.1.2 | EAAR | 419/420 | ed | The table entry regarding PMCF/PMPF has a missinga full-stop. Please change to “PMCF/PMPF plan may be required. These could include:”. | The table entry regarding PMCF/PMPF has a missing full-stop. Please change to “PMCF/PMPF plan may be required. These could include:”. | accept
5.1.2 | EAAR | 419/420 | ed | Change “annex XIII” to “Annex XIII” in the last paragraph. | Change “annex XIII” to “Annex XIII” in the last paragraph. | accept
figure 437 | EAAR | 437 | Ed | Typo in Figure 2: PMCF/PMF studies | PMCF/PMPF studies | accept
5.4 | EAAR | 467 | Ed | “… as explained in chapter 6” --> reference to wrong chapter |  | accept
5.4 | EAAR | 479 | ed | Change full-stop to comma to get “in a summary report, the PMS report or”. | Change full-stop to comma to get “in a summary report, the PMS report or”. | accept
5.4 | EAAR | 480 | ed | Footnote 8 is not shown. There is only an 8 at the end of the sentence. | Add footnote 8. | accept
5.4 | EAAR | 482ff | ed | Sentence should start with capital “T” and the formatting of the whole paragraph should be reviewed. There are three full-stops with non-capital letters afterwards. Some may be commas? | Review and revise paragraph. | accept
6 | EAAR | 498 | ed | Change “form” to “from”. | Change “form” to “from”. | accept
6 | EAAR | 503 | ed | Add full-stop at the end of the sentence. | Add full-stop at the end of the sentence. | accept
annex I | EAAR | 530 | Ge | For Class IIa/b MDR devices and Class C IVDR devices, the PSUR needs to be “made available to NB” – an explanation of what this specifically means is lacking |  | reject. Not in scope of this document.
annex I | EAAR | Page 19 | ed | In the PMS Plan section remove space in front of “Describes”. | In the PMS Plan section remove space in front of “Describes”. | accept
ge | EFPIA | N/A | ed | Use of abbreviations and full wording for a given term is not consistent throughout the document. e.g., both “PMS system” and “post-market surveillance system” are both used in-paragraph with no apparent consistency. | Check document for consistency of terminology and use of abbreviations. | accept
1 | EFPIA | 123-125 | Ge | It would be helpful to clarify which report is being referred to. Is this the PSUR or PMS report? |  | both reports therefore only report is mentioned
1 | EFPIA | 86-96 | ge | This section doesn’t mention if device parts of single integral products are in or out of scope | Specify that device parts of single integral products are out of scope | reject - refer to device definition in intro for inclusion criteria
3 | EFPIA | 165-168 | ge | There is no definition of “real world” information in the EU MDR. This also applies in section 5.2, where the term is used again. | Remove the term “real world”. | accept
PMS as an integral part o | EFPIA | 212-213 | ge | This bullet states that a manufacturer must identify potential sources of PMS data and the amount of data to be collected. Use of “must” implies the actions mandatory and that the language used is consistent with the corresponding regulation. While the description of the PMS Plan within the Medical Device Regulation covers identifying data sources to be collected (although not explicitly), it does not mandate describing the amount of data to be collected. | Re-word the bullet in such a way that defining the amount of data to be collected is not a mandatory requirement for the manufacturer. If appropriate, make this an optional activity instead. | Accept - remove the "amount of data" (not in the regs)
PMS as an integral part o | EFPIA | 207-208 | ed | This text is formatted as a heading, but is inconsistent with the rest of the document | If intended as a heading, number as section 3.1 | accept
PMS as an integral part o | EFPIA | 214-215 | ed | This bullet focuses only on characterisation of performance; however, the data collected in the first bullet is also used in the characterisation of safety of the device. | Must collect (…) to allow a correct characterisation of the safety and performance of the device. | accept
PMS as an integral part o | EFPIA | 228-231 | ed | This bullet describes the purpose of the PMS plan. This doesn’t fit with the content of the rest of the list, which details actions required of the manufacturer. | Include a requirement for the PMS plan as the first bullet, aligned to ISO/TR20416:2020, e.g.: must plan the PMS activities to be conducted before the first placing on the marketing of the medical device and update as necessary during the product life cycle The bullet which describes the purpose of the PMS plan can be separated from the bulleted list and placed in a regular paragraph, starting “the PMS plan referred to in the first bullet above shall address the collection…” | accept
4 | EFPIA | 243-396 | ed | Format of sub-headings throughout section 4 is inconsistent with the rest of the document | Re-number a) through h) as 4.1 through 4.8 | accept
grafic 1 | EFPIA | 241 | ge | Update “benefit-risk determination” and “improve risk management” are grouped together in a single part of the diagram. | Consider separating benefit-risk determination from improve risk management as while the former is covered under risk management, it is also covered under clinical evaluation and post market surveillance | agreed, graphic will be taken out or replaced with a different graphical representation.
grafic 1 | EFPIA | 241 | ed | Diagram is aligned to Article 83.3, but does not reference this article. | Add reference to Art. 83.3, e.g. as a footnote | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | EFPIA | 243-269 | ge | Interface with the clinical evaluation for the verification of the benefit-risk determination is not mentioned | As suggested for the diagram, have a separate section for benefit-risk determination which covers how PMS is used to update the benefit-risk determination and the interface with clinical evaluation and risk management. | agreed. Text will be aligned once the graphic repesentation is finalized.
4a | EFPIA | 253 | ge | Unclear why this first bullet point is relevant here as this is quite general and it is about the PMS activities in general including the other points below-mentioned from b to h . | Remove bullet point and consider including in the bulleted list in the unnumbered section within section 3 | accepted.
4c | EFPIA | 333 | ge | It is unclear exactly what is being referred to in “the risk management” | Specify what this refers to – i.e. is it the risk management: - plan - activities - system - file… | Reject. The text aligns with the wording in the regulations.
4c | EFPIA | 299 | ge | Bulleted list, which references Annex XIV Part B, is missing the bullet about ensuring acceptability of the clinical evidence and of the benefit-risk ratio | Add an additional bullet point: “ensuring acceptability of the clinical evidence and of the benefit/risk ratio” | accept
4c | EFPIA | 314 | ge | Not all points from the bulleted list in Annex XIII, Part B are included | Add the missing item regarding “identifying possible systematic misuse” for consistency with the regulation | Reject, already in the text
4c | EFPIA | 288 | ed | Missing full stop. | Insert full stop between ‘”clinical data”’ and ‘Annex XIV’ | accept, has been rephrased.
4c | EFPIA | 330 | ed | Sentence ends “clinical evaluation r/” | Correct to “clinical evaluation report/” | Reject - sentence has been restrcutured.
4d | EFPIA | 350-351 | ed | Incorrect report name, missing capitalization and full stop. | Correct “Periodic Summary Update report The post-market…” to “Periodic Safety Update Report. The post-market” | accept
4f | EFPIA | 374 | ge | Clarify what is meant by the “use” of the medical device, in particular to differentiate usability vs indication for use. | Assuming this refers to usability, propose amendment such as “to improve the instructions for use and design features affecting usability of the medical device”. | partially accepted - text modified
4g | EFPIA | 376-385 | ge | This section is not clear on the expectations of the manufacturer. Is the manufacturer expected to maintain an up-to-date list of similar devices and in what type of report/document should the manufacturer present an evaluation of consideration for similar devices, considering some aspects of a competitor device design may be unknown to the manufacturer? | Provide clear, defined examples of how to communicate and evaluate the risk to similar devices. | accept paragraphs are rephrased
5 | EFPIA | 403 | ed | Paragraph 6.1.1 does not exist. | Correct to 5.1.1 | accept
5 | EFPIA | 405 | ed | Paragraph 6.1.2 does not exist | Correct to 5.1.2 | accept
5.1 | EFPIA | 412 | ge | There is no guidance in the planning section about until when a PMS plan is needed. | Add a sentence at the end of this section stating that “the PMS plan is maintained throughout the lifetime of the medical device” | accepted
5.1.1 | EFPIA | 413 | ge | The information included in column “how to collect it” is currently not consistent across the different topics and is not adding much value. | Provide practical suggestions across all entries on how a manufacturer can collect data – e.g. for feedback, complaints and incidents providing other methods such as a dedicated support line or email address for users to contact and providing information in the IFU in addition to the proactive method of using surveys; suggestions on how to utilize registers to collect PMS data | reject, it is beyond this guidance to give specific how, more the general guidance
5.1.1 | EFPIA | 413 | ge | The column “how to utilize it” could provide more practical examples. | Align the column “how to utilize it” with ISO /TR20416:2020 – Annex A as it provides details on why the information on the different data sources are useful for. | reject, however, text will be more clear why all information is necassery
5.1.1 | EFPIA | 413 | ge | The table in section 5.1.1 does not cover guidance on what social media data a manufacturer should collect and how to evaluate that information | Add social media as potential data source in the section on feedback and complaints and align the guidance on collecting and utilizing this data with ISO/TR20416:2020. | reject
5.1.1 | EFPIA | 414 | ge | “Information concerning serious incidents, including information from PSURs and field safety corrective actions” Trending is already covered in its own separate section | Remove references to trending from this section and add to the relevant section in the table; otherwise add trending considerations for other data sources such as feedback and complaints where this activity is required. | acepted, text reworded
5.1.1 | EFPIA | 414 | ge | Publicly available information about similar medical devices “how to utilize it” column – it does not mention how to use the information posted by NCA on similar medical device technologies. | Add in the “how to utilize it” column that the evaluation from NCA on similar medical technologies could help the manufacturer with the state of the art for their product. | reject - evaluations of NCAs are not public
5.1.1 | EFPIA | 414 | ed | There are two rows in the table titled “Information, including feedback and complaints, provided by users, distributors and importers”, which contain similar, but not identical, information on how to collect/utilize | Align on correct language and remove redundant entry appearing later in the table. Additionally, include text to encourage reporters to provide as much detail as possible, as seen in the sections on serious and non-serious incidents | accept
5.1.1 | EFPIA | 414 | ed | “Relevant specialist or technical literature, databases and/or registers” There is no specification on which type of devices we are referring in the statement “Ifac new risks are identified in relation to other devices (…)” | Clarify that this concerns to similar devices. Suggestion: If new risks are identified in relation to other similar devices (…)” | accept
5.1.2 | EFPIA | Page 12, Information from trend reporting | 2 | Our understanding of MDR article 88 and IVDR article 83, is that it states that Manufacturers are only required to report those trends (i.e. statistically significant increases in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side-effects) via the electronic system that: Are “non-serious incidents” or, Are “expected undesirable side-effects that could have a significant impact on the benefit-risk analysis referred to in Sections 1 and 8 of Annex I and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits.” In the text stated in the table, it seems like any statistically significant increase in undesirable side effects are to be reported regardless of whether it is assessed to have impact on benefit-risk analysis or not. We suggest that it is clarified in this section that in order to qualify for reporting via the Electronic system mentioned in EU MDR article 88 and EU IVDR article 83 trends identified for expected undesirable side effects should be assessed to have a potential significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. In addition, we suggest that a reference is inserted for the Q&A on trend reporting and that the abbreviation GSPR is explained | We suggest changing the wording to: “Articles 88 MDR and article 83 IVDR require manufacturers to monitor non-serious incidents and undesirable side effects to identify any statistically significant increase in the frequency or severity of non-serious incidents and undesirable side effects. The Manufacturer shall submit a trend report for any statistically significant increase in the frequency or severity of: Non-serious incidents and expected undesirable side-effects that could have a significant impact on the benefit-risk analysis in relation to the GSPRs* and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. The outputs from non-serious incidents including data on undesirable side effects should be gathered and consideration should be given to whether any trends are evident. See further the Q&A on trend reporting *General Safety and Performance Requirements | Trending TF will be consulted.
5.1.2 | EFPIA | 419 | ge | — effective and appropriate methods and tools to investigate complaints and analyse market-related experience collected in the field; The current explanation focuses on the “what”, but does not provide clarification on the “why” or the “how”. | Suggest to state in the explanation column that manufacturers should establish processes about choosing an appropriate method for data analysis depending on the quantity and quality of collected data and how it should be used. | tabel is revised
5.1.2 | EFPIA | 419 | ge | methods and protocols to manage the incidents subject to the trend report as provided for in Article 88 MDR and article 83 IVDR, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period;- the first sentence is a repetition of what a PMS plan should describe. The MDCG Q&A trend is not yet published. | Amend this sentence to cover all applicable aspects of trend reporting, as this extends beyond complaints and market-related experience (e.g. serious and non-serious incidents) Ensure that the MDCG Q&A trend document is published before or at the same time that the MDCG for PMS. | tabel is revised
5.2 | EFPIA | 437 | ge | Social media is missing as a source under User Feedback Audits/inspections are mentioned as a data source for the first time in the document in this diagram, but are not elaborated on at all elsewhere in the document | Include social media as a source under User Feedback. Provide guidance in section “The post market surveillance plan shall address the collection and utilization of available information” on how to collect social media data and audits and inspection information. | reject is already in the figure
5.2 | EFPIA | 430 | ge | acc | Remove reactive or passive adjective from the sentence | accept
5.3 | EFPIA | 450-456 | ge | This paragraph is not under the scope of the assessment and analysis of data. The PMSR or PSUR are the reports that summarise the data analysis and draw also the conclusions of the PMS plan. | Relocate to section 5.4. Instead include a paragraph explaining that data is used to determine new risks and changes to existing risks against the risk management file. | reject, it is in th eright place
5.1.1 | Eurom VI | 414 / page 11 & 13 | ge | “Information, including feedbacks and complaints, provided by users, distributors and importers”, paragraph is doubled (page 11 & page 13). | This can be summarized in one paragraph. | accept
5.1.1 | Eurom VI | 414 / page 12 | ge | Paragraph “Information from trend reporting” & “Records referring to non-serious incidents and data on any undesirable side-effects” are content-wise the same (page 12). | This can be summarized in one paragraph. | accept
5.1.2 | Eurom VI | 419 / page 15 | ge | The row which describes “systematic procedures to identify and initiate appropriate measures including corrective actions;” is difficult to understand. | This sentence can be simplified. | accept - sentence can not be changed as this is the wording from the MDR. However, the text in the explanation has been updated.
ge | CA5 |  | ge | “The post-market surveillance system referred to in Article 83 shall be based on a post-market surveillance plan” consequently the PMS Plan is the key document. Firstly, the PMSP should be defined, described, the MF’s responsibilities should be identified by device and emphasis that it is part of the device technical documentation. Important to connect the PMSP to the post-market surveillance report/PSUR, since on those reports the PMS data is reported and the PMSP is updated based on the received results. Also include the difference on the data analysis; on the PSUR only 4 last years are reported but within the PMS activities a longer period of data is used. Secondly describe the connection to the other required documents and their interaction. | I consider that the figure “Data source feeds to be used for generation of reports” removed from the PSUR guidance could be useful in the PMS guidance. I consider the figure shows the continuous process interaction of different sources. Add the references adequately. So, the structure of the document could be clearer, for example in the section 4 instead of using letters use numbers. In addition, I am not sure whether using a list from article 83 is the best way of data presentation. The text contains plenty of regulation text and less advice, what the regulation indicates in practice. | accept, figures are all taken out
intro | CA5 | 63-67 | ge | I prefer to change the order of the paragraph so that it beginnings with the new requirements of the regulations. In addition, the guidance document is intended for manufacturers and the part concerning CAs is not necessary. | The Medical Devices Regulation (MDR) and In-Vitro Diagnostic Medical Devices Regulations (IVDR) laid out substantially more requirements to the manufacturers to perform the post-market surveillance activities compare to the former legislation. | Removed NCA oversight requirement as this is Market Surveillance, not PMS
intro | CA5 | 68-71 | te | I consider the first sentence of the paragraph could be removed and instead add text, what the post-market surveillance e.g., definition for it or an overview about it. May be the two last paragraphs could be unified. | The following is just an example: PMS is the manufacturer’s process to collect and analyze data to ensure the device safety, performance, and its benefits within the entire life cycle of the device.
1 | CA5 | 86-92 | te | Based on the Trend report Q&A document (Q4) it is possible to use the regulation process also for the old devices so the issue should be mentioned in this context. |  | trend tf will be consulted
3 | CA5 | 137-205 section 3 | ge | In different paragraphs repeating the same issues, so the section should be completely rewritten and remove the unnecessary repetition. Furthermore, no reference to the Annex III or PSUR, which are the basis for the PMS activities. In addition, no mention that the device should reach the benefits intended. |  | accept the need for revision
3 | CA5 | 141-142 | te | Add the responsible actors, may be the text could be relocated to the section 2 Terms and definitions | These surveillance actives need to be distinguished between “market surveillance” activities by the competent authorities and “post-market surveillance” activities by the manufacturer. | delete ca activities
3 | CA5 | 165-166 | te | I prefer the term compare in this connection. I consider that the new data should also be compared the previously collected PMS data not only to the pre-market data. | The PMS system is intended to continuously monitor compare the results of the pre-market phase (also known as development phase) to the with “real world” information data from post-market | accept
3 | CA5 | 155-158 | ed | The whole paragraph could be dislocated to the instruction to explain what the PMS is. |  | accept
PMS as an integral part o | CA5 | 214-215 | te | An ambiguous sentence, should be clarified | Perhaps list the required datasets or make a reference to the PSUR guidance. In addition include the manufacturer’s responsibility to identify the used datasets by device and the monitoring interval. | make more practical
PMS as an integral part o | CA5 | 216-217 | te | The bullet point is included in the first bullet point | See the comment above | make more practical
PMS as an integral part o | CA5 | 219-220 | te | Instead of using the parentheses us the bullet points. Prefer to include methods to | · must establish effective and appropriate methods and processes to assess the collected data for example define by device The analytical methods The trending period Indicators and threshold Scheduling the planned actions Triggers for preventive and corrective actions | make more practical
PMS as an integral part o | CA5 | 223-227 |  | The sentence does not bring any new idea or clarifies the issue, so it could be deleted | Delete the sentence | accept text revised
PMS as an integral part o | CA5 | 228-231 |  | The sentence could on the first paragraph of this section | Transfer the sentence in the bigging of the section | accept text revised
4a | CA5 | 254 | te | Reference to PSUR and Post-market surveillance report is missing | Suggest to add the references, since those are the PMS reports, or the information could be provided on section 3 | accepted, will be explained in the main text of the guidance in Section 2 of the document
4a | CA5 | 262 | te | The thresholds i.e., acceptance rate relates to all risks. | Update the text to cover the coherent whole. | rejected - it is unclear what is meant.
4c | CA5 | 311 |  | TYPO: • dentifying and analysing emergent risks on the basis of factual evidence | • identifying and analysing emergent risks on the basis of factual evidence | accept
5.1.1 | CA5 | 414 | te | Information, including feedback and complaints, provided by users, distributors and importers; was twice on the table, row 2 and 7. The text was not identical | Remove the repeated text. The table includes lot of text may be usage of the lists will make it easier to read. In addition, the text should be restricted to the column headings How to collect data and how to utilize it | accept
5.1.1 | CA5 | 414 table, last row |  | “The data that is gathered via the activities listed within Annex III 1.1(a) should be used to monitor any emerging risks as well as to help demonstrate continued acceptability of the performance and safety of the device. “ | The text is not related solely to the name of the row Publicly available information about similar medical devices nor the usage of the data. | accept
5.1.1 | CA5 | 414 table, row 6 |  | “Articles 88 MDR and article 83 IVDR require manufacturers to trend and report any statistically significant increase in the frequency or severity of non-serious incidents, or undesirable side effects. This includes those that could have a significant impact on the benefit-risk analysis in relation to the GSPRs and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits.” data on undesirable side effects “The manufacturer should consider whether field safety corrective actions are required to be issued including field safety notices”. | I suggest to remove the text, the MTR should be handled previously in the text part. data on undesirable expected side effects. The manufacturer should define the trending period, methods to detect trend on the PMSP. In general, the CAPA follows the MTR, also the FSCA is possible. | accept - TF on Trend Reporting will be consulted to ensure alignment
5.1.1 | CA5 | 414 table, row 6 |  | “This information should be used to help determine whether there are any new or emerging risks associated with the device.” | The change in the severity and rate of the known risks should also be considered and to detect the specific patient groups which might be exposure to higher risk than other patient groups. | TF on Trend Reporting will be consulted to ensure alignment
5.1.1 | CA5 | 414, table row 3 |  | “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. This information should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR.” ….” should result in an update of the periodic safety update report (PSUR) in line with the PSUR guidance” | Trend reporting is related to the incidents, expected side-effects and erroneous results, but not to the serious incidents. In addition, the article 89/84 concerns the serious incidents. I consider that the text on Annex III including information from PSURs refers to the already published PSURs, to utilize the previously collected and analyzed data, not to start an PSUR update process which is performed according to the PSUR guidance. | accept
5.1.1 | CA5 | 414, table row 4 |  | “results compared with the output of the risk management process to identify any new or emerging risks” | The change in the severity and rate of the known risks should also be considered. | See also coment in 504. TF on Trend Reporting will be consulted to ensure alignment
5.1.2 | CA5 | 418, table 2 |  | “5.1.2. The post-market surveillance plan shall cover at least:” | I am not sure if the table adds information, at least the table should be clarified and shortened | accept, table has been reviewed and content modified.
5.2 | CA5 | 422 | te | 5.2. Collecting necessary data | Suggest to change the heading Data collection | accept
5.2 | CA5 | 423 | te | Developing a proactive PMS approach Is this a heading or a subheading? | All headings should be numbered logically. In addition, to which heading the text is related. | accept
5.2 | CA5 | 424-428 | te | Reference is missing | Add the reference | accept
5.2 | CA5 | 431-433 | te | The regulation refers to the proactively collected data and the real-world data is one part of proactive data. | Please add the reference where the real-world data is required. | accept
5.2 | CA5 | 433-434 | te | In some instances, manufacturers may be required to generate new data relating to real-world use of the device. | Please add some examples and the reference. | accept
5.3 | CA5 | 443-445 | te | The PMSP should cover the methods and processes for all collected data, it is not necessary to raise the complaints. In addition, the timing of the processes is missing. | Update the sentence to cover all data and add the timing for the actions to be taken | reject, Timing is not for the pms plan
5.3 | CA5 | 441 | te | Overall Utilization of the Gathered Data | All headings should be numbered logically. I am not sure, whether the data analysis is related to the data utilization. Firstly, you analyze the data, secondly draw the conclusions, and thirdly utilize the received results | accept
5.4 | CA5 | 479-480 |  | Please, check, some unfinished sentences. the PMS report or the PSUR8. |  | accept
6 | CA5 | 494-496 |  | Please, check: implement appropriate means to apply any necessary corrective action, In that context, it shall report in accordance with Article 87(1) to the competent authorities any serious incidents or field safety corrective actions or both as soon as it learns of them. | implement appropriate means to apply any necessary corrective action. In that context, it shall report in accordance with Article 87(1) to the competent authorities any serious incidents or field safety corrective actions or both as soon as it learns of them. | accept
6 | CA5 | 504-507 |  | Please, check: The PSUR is a summary of the information collected and analysed, together with the conclusions and the action(s) taken. the PSUR and the other information gained in the post-production phase is part of the documentation for custom-made devices (see MDCG 2021-3) according to Section 2 of Annex XIII of the MDR | The PSUR is a summary of the information collected and analysed, together with the conclusions and the action(s) taken. The PSUR and the other information gained in the post-production phase is part of the documentation for custom-made devices (see MDCG 2021-3) according to Section 2 of Annex XIII of the MDR. | accept
intro | FMoH | 74-77 | ge | The subparagraph do not cover the objective of PMS (updating the MDs technical documentation). | Once placed on the market, European regulations oblige the manufacturer to perform an active role during the post-market surveillance phase for the medical device and to monitor whether the medical device continues to meet the general safety and performance requirements specified in the regulation by systematically and (pro)actively gathering information from post-market experience with their medical device in order to identify and manage risks that reveal in the post-market phase. | Accept - added wording about 'identify any need for preventive and corrective action'
1 | FMoH | 86-87 | ge | Move the description of covered products from section 2. | This guidance covers medical devices, accessories for medical devices, products listed in Annex XVI of the MDR, in vitro diagnostic medical devices (IVDs), accessories for IVDs and custom-made devices. | accept - was moved to introduction
1 | FMoH | 102-104 | ge | The sentence do not fit to the previous sentence. The requirements for the QMS do not ensure the implementation of the PMS. The implementation of the PMS also depends on the PMS plan. | Delete the sentence. | reject - art 10 states that PMS is addressed by the QMS and points to art 83 (PMS Plan)
1 | FMoH | 121-122 | te | The first sentence of the subparagraph is a repetition of the previous subparagraph. | Delete the sentence. | partly accept
1 | FMoH | 122-123 | ge | A deeper description of PMS plan is missing. | Include the sentences: The PMS system based on a PMS plan referred in Article 84 MDR / Article 79 IVDR. For devices other than custom-made devices, the PMS plan shall be part of the technical documentation specified in Annex II. Update the sentence in line 123: Both regulations also describe that a report […] | reject is explained later in th edocument
1 | FMoH | 80-81 | ed | The PMS plan described in section 5.1 is missing. | To describe the PMS system and processes required by Article 83 and 84 MDR / Article 78 and 79 IVDR | accept
1 | FMoH | 100 | ed |  | Delete the commas in the references. | accept
2 | FMoH | 126-135 | ge | Section 2 is not needed. Line 130-133 is moved to section 1. | Delete section 2. | changed to subsection
3 | FMoH | 161-163 | te | Clarify the PMS’s objective. | […] performance requirements as well as acceptable benefit-risk-ratio, the manufacturer needs to (pro)actively gather information from post-market experience. It is vital to understand that a manufacturer needs to demonstrate how they observe and check the safety, performance and benefit-risk-ratio of the medical […] | accept
3 | FMoH | 174 | te | Clarification that collection and evaluating clinical data/performance data from the use of the device. | Add at the end of the sentence: manufacturer to update the clinical evaluation (PMCF)/ performance evaluation (PMPF). | accept the need for revision
3 | FMoH | 177 | ge | Evaluation of undesirable side-effects/interference and cross-reaction and the acceptability of the benefit-risk-ratio is also the objective of the PMCF and PMPF. | Add at the end of the sentence: and the evaluation of the undesirable side-effects (MDR)/ interference(s) and cross-reaction(s) (IVDR) and of the acceptability of the benefit-risk-ratio. | accept the need for revision
3 | FMoH | 184-185 | ge | Unclear what is meant. It is not the obligation to monitor information. | Delete “and thereby demonstrate to actively monitor information”. | accept the need for revision
3 | FMoH | 155-158 | te | PMS include activities carried out by the manufacturer to ensure safety and performance as well as acceptable benefit-risk-ratio. Quality requirements are not part of PMS. | In simpler terms, PMS includes the systematic, planned activities carried out by a manufacturer to ensure the safety and performance as well as acceptable benefit-risk-ratio of a medical device. These activities include the collection and evaluation of information to the safety, performance as well as risks and benefit of a medical device once it has been placed in the market. | accept
3 | FMoH | 170 | ge | The whole PMCF and PMPF is addressed in the PMS. | Update reference: Annex XIV, Part B of the MDR/Annex XIII, Part B of the IVDR | accept
3 | FMoH | 171 | te | PMCF and PMPF is an obligation. | […] emphasize the obligation of a manufacturer […] | accept
3 | FMoH | 171 | te | To monitor is not the obligation of the PMCF and PMPF. | manufacturer to update the clinical evaluation (PMCF)/ performance evaluation (PMPF). | accept
3 | FMoH | 198 | te | Clarify the PMS’s objective. | […] is to provide a holistic view of the safety and performance as well as acceptability of the benefit-risk-ratio of the medical device […] | accept
PMS as an integral part o | FMoH | 228-231 |  | It is not clear what is meant. Also did not fit to “This means that a manufacturer”. | Rewrite the bullet point. | accept text revised
PMS as an integral part o | FMoH | 212 | te | The amount of data imply that information gathering has an end. | Delete “and the amount of data” | accept
5 | FMoH | 403 + 405 | ed | Section 6.1.1 and 6.1.2 does not exist. | Update references: Section 5.1.1/5.1.2. | accept
5.3 | FMoH | 462 | ed |  | Update references: Annex III 1 b | accept
5.4 | FMoH | 482-485 | ge | What is meant? | Rewrite subparagraph. | accept
3 | CA10 | 170 | editorial | Original text: “Annex XIV, Part B, Sections 5 and 6 MDR/Annex VIII, Part B, Sections 4 and 5 of the IVDR” It seems to be a typing error. The Annex VIII of the IVDR is the Classification Rules; while the Annex XIII part B is about the Post-market Performance Follow-up, which corresponds to the subject of this paragraph. | “Annex XIV, Part B, Sections 5 and 6 MDR/Annex XIII, Part B, Sections 4 and 5 of the IVDR” | accept
4a | CA10 | 264 | general | Original text: “(as referred to in Section 3 of Annex I MDR)” The IVDR is not mentioned. | “(as referred to in Section 3 of Annex I MDR / IVDR)” | accept
5 | CA10 | 403 | editorial | Original text: “ In paragraph 6.1.1 you can find an explanation of each indent of Annex III part 1(a).” There is no paragraph 6.1.1 in this document, the above-mentioned content correspond to the paragraph 5.1.1. | “ In paragraph 5.1.1 you can find an explanation of each indent of Annex III part 1(a).” | accept
5 | CA10 | 405 | editorial | Original text: “ Paragraph 6.1.2 explains the content of Annex III part 1(b).” There is no paragraph 6.1.2 in this document, the above-mentioned content correspond to the paragraph 5.1.2. | “ Paragraph 5.1.2 explains the content of Annex III part 1(b).” | accept
ge | MedTech Europe | n/a | ge | What is the purpose of this guidance? Is it meant to offer more guidance than existing documents around HOW we perform some of the analysis of the PMS data and how it links to the risk file? The current document does not seem to offer anything beyond what we already can extrapolate from the MDR and guidance that already exists.
ge | MedTech Europe | 23 | te | Line 23 – should MDR be included here? | Change to: “ The PMS System as an integral part of the manufacturer's quality management system referred to in MDR Article 10(9) (IVDR 10(8))” | accept
ge | MedTech Europe | n/a | ge | MDR requires to reduce risks „as far as possible“, often in combination with „designed and manufactured“ (see MDR, Annex I, Chapter I, several clauses) | The guidance should provide more clearance for manufacturers. Is this applicable for the PMS phase as well or is this purely a requirement for the design and manufacturing phase? It is almost impossible to reduce all risks as far as possible, if the technology has improved after placing the device on the market. This can be done for new versions of the device, but not using the recall process. | accept
ge | MedTech Europe | n/a | ge | The guidance seems to be highly repetitive. The document appears to be an early draft of a document that needs significant streamlining and simplification. | Revise entire text of the guidance with the goal of simplifying the text and have no repetitive text. | accept
ge | MedTech Europe | n/a | ge | Concentrate the document on the most relevant points to ensure that this document adds benefit and is providing meaningful guidance to the intended target group |  | accept
ge | MedTech Europe | n/a | ge | A number of prior MDCG guidance documents provide templates to be used to meet MDR requirements for PMS reporting (PSUR/PMCF). Something needed for IVDR. | Please consider providing PMSR/PSUR/PMPF/PMCF templates manufacturers can use to meet the requirements for both IVDR and MDR reporting | accept
ge | MedTech Europe | n/a | ge | If the guidance focus is on MDR/IVDR products only that is clear, if not clearly stated MDR/IVDR devices | The answer to this requirement could be: “PMS also needs to be done for legacy devices”. | accept
ge | MedTech Europe | n/a | ed | Language may be clearer and less narrative. In some instances, the text could be shortened significantly while becoming more precise and clearer. |  | accept
ge | MedTech Europe | n/a | ed | Decompose the lengthy texts and break it down into facts |  | accept
ge | MedTech Europe | n/a | ed | Review the document for ambiguous terms and either remove or define them in a better manner, for example (not limited to): “registry studies”, “expert use groups”, “focus group”, “post market clinicals” |  | accept
ge | MedTech Europe | n/a | ed | Determine when a legal reference should be part of the main text or when it would be better to place it as a reference (footnote) to make the document more readable. |  | accept
ge | MedTech Europe | n/a | ed | Review the figures and determine what the aim of the figures is; the first figure could also be simply presented as a bullet list. Regarding Figure 2, the aim of the figure is unclear as the structure and systematic is not clear to the reader. I don’t believe we should add figures to a document just for the sake of having a figure if there is no additional value and benefit in understanding to the reader. |  | accept
ge | MedTech Europe | 23 | ed | Line 23 – should MDR be included here? | Change to: “ The PMS System as an integral part of the manufacturer's quality management system referred to in MDR Article 10(9) (IVDR 10(8))” | accept
1 | MedTech Europe | 111-114 | Ge | It is not a sentence. There is something missing to add sense to the paragraph. | Clarification is needed to understand better the paragraph | accept
1 | MedTech Europe | 111 | Ge | Line 111 is written “…for the risk management to be found in section 3 of Annex I” - section 3 Annex I refers to: CHAPTER III REQUIREMENTS REGARDING INFORMATION SUPPLIED WITH THE DEVICE | Check if Chapter II is meant | reject section 3 is not chapter 3
1 | MedTech Europe | 120 (and throughout) | Te | The guidance (and the regulation itself) is not clear on the difference between a CAPA corrective or preventative action versus a Field Safety Corrective Action. In some places, it appears to refer more to a FSCA versus a CAPA while other places appear to mean both. | It would be helpful in the guidance to outline what the regulation means in terms of corrective actions (FSCAs vs CAPAs) For example in line 124 it is mentioned “rationale and description of any preventative and corrective actions taken. The use of the term rationale appears to point more to a FSCA rather than a CAPA.
1 | MedTech Europe | 87-92 | Ed | Consider including note for custom made devices. Consider including statement that the PMS requirements are required throughout the device lifetime. | Change to: “The PMS requirements in the MDR and the IVDR apply to MDR/IVDR devices as well as “legacy devices” and custom made devices certified under the Medical Device Directive (MDD) (93/42/EEC), the Active Implantable Medical Directive (AIMDD) (90/385/EEC) or the In Vitro Medical Directive (IVDD) (98/97/EEC) according to Article 120(3) MDR and Article 110(3) IVDR throughout the entire device lifetime. Requirements for IVD Legacy devices have been detailed in MDCG 2022- 8.” | partly accepted
1 | MedTech Europe | 99 | ed | Eliminate redundancy from the text | Delete: “Article 10 of the” | accept
1 | MedTech Europe | 101 | ed | Alternative wording replacing “In particular” with “Specifically” | Change to “Specifically, Article 10, (9) MDR and Article” | resolved/ sentence was removed
1 | MedTech Europe | 101 | ed | Article missing in Line 101: “IVDR require manufacturer” | Change to: “IVDR require a manufacturer” | resolved - plural used/ decline see line 24 AT solution
1 | MedTech Europe | 104 | ed | Preposition missing in Line 104: “risk class and appropriate the type of device.” | Change to: “risk class and appropriate to the type of device.” | reject (see above)
1 | MedTech Europe | 104/5 and 115/6 | ed | The information is somehow double | The information Could be combined with line 104/5 and 115/6 | see comment on line 73 above
1 | MedTech Europe | 117 | ed | Delete “be suitable to” to make the text clearer | Change to: “post-market surveillance system shall actively and systematically gather,” | decline is the text from art 83(2)
2 | MedTech Europe | 126 | ge | Some definitions are missing: proactive or active vs passive like in CEN ISO/TR 20416:2020, Post-market surveillance for manufacturers; Real world data and real world evidence (like FDA); general and specific methods and procedures of PMCF is not clearly explained; | The definitions should be incorporated. Define proactive and active in terms of data collection and monitoring as we see variation in interpretation from Notified Bodies (i.e. NB not considering literature review active) | REJECT adding definition - not the source uis active but the action taken by the manufacturer. Definitions from not legal documenst we can not copy.
2 | MedTech Europe | 129 | ed | Alternative wording: delete “yet” and replace with “previously” | Change to: “not previously explained in existing MDCG guidance’s the following definitions apply” | accept
2 | MedTech Europe | 130-133 | ed | Medical Device definition provided in MDR Art 1(4) and 2(1) | Use the definition provided in MDR/IVDR | accept
3 | MedTech Europe | 178-182 | ge | There is no consensus between the Regulations, MDCG guidance documents, and notified bodies on what is considered general vs. specific activities. Surveys are considered specific PMCF activities by TUV and BSI. Additionally, they are alluded to being specific activities in the MDCG 2020-7 PMCF plan guidance (list of PMCF examples in section C). The language specifying general vs. specific should be removed unless it accurately reflects the accepted methods by notified bodies as understood by the requirements laid out in the regulations. | Remove “gathering of clinical experience (e.g. via general surveys)” as surveys may be considered specific methodology. | accept the need for revision
3 | MedTech Europe | 199 | Te | Needs clarification. It is allowed under MDCG 2020-7 to transition from specific activities to general but can that still be construed as actively monitoring? Example, device registry needs to pre-specify the quantity – once that quantity is achieved, registry monitoring can stop (Section C, page 7 of pdf of 2020-7). Similar for conducting a post-market CI; eventually it will end. |  | accept the need for revision
3 | MedTech Europe | 143 | ge | Is it expected that Market surveillance is continued during the life of the product to help conclude if the product still remains in state of the art? Is this pre-commercial activities? Clinical trials? What are those activities? | Clarification is needed on the scope of Market Surveillance. | part regarding CA will be deleted not for guidance
3 | MedTech Europe | 155-158 | Ge | Consider including statement that the PMS requirements are required throughout the device lifetime. | PMS includes the systematic, planned activities carried out by a manufacturer to monitor the safety, quality and performance of a medical device throughout the entire device lifetime. | accept
3 | MedTech Europe | 180 | te | Not entirely clear how is such a “general survey” defined | Clarification is needed on what is meant by “general survey”. Suggestion: to delete the term “general” | accept
3 | MedTech Europe | 185, Footer page 5 | te | Footer page 5 (80) MDR seem to be missing IVDR information | If applicable, reference to applicable IVDR articles should be added | accept
3 | MedTech Europe | 134 | ed | Using different choice of words | Change to: “We will provide if necessary, a definition in the footnote or a reference” | deleted
3 | MedTech Europe | 134 | Ed | All definitions should be collected in this section. Foot notes may be used in addition if helpful/relevant |  | reject
3 | MedTech Europe | 138 | ed | Alternative wording suggested | Change to: “Surveillance activities begin when the manufacturer of a medical…” | accept
3 | MedTech Europe | 147 | ed | The text in Line 157 is not entirely clear. | Revise text to make it clearer. | accept
3 | MedTech Europe | 149-154, 155-158, 155-156, 157-158 | ed | Text in those paragraphs is not entirely clear and sometimes redundant | Simplify the text and reduce redundancy. Example: lines 149-154 and lines 155-158. Additionally, lines 155-156 and 157-158 are essentially identical in information. | accept
3 | MedTech Europe | 159 – 164 | ed | The text in Line 159 – 164 seems repetitive and could be simplified. | The text could be simplified and the repetition reduced | accept
3 | MedTech Europe | 157 | ed | Minor editorial change: delete “to” and replace with “on” | Change to: “activities include the collection and evaluation of information on the safety, quality, and” | accept
3 | MedTech Europe | 158 | Ed | Grammar correction needed: … it has been placed in the market… | Change to: … it has been placed on the market… | accept
3 | MedTech Europe | 165 | ed | The line seems to be redundant to previous paragraphs | This paragraph may require attention. The language is not entirely clear and intent may be misunderstood. | accept
3 | MedTech Europe | 170/171 | Ed | …/Annex VIII, Part B, Sections 4 and 5 of the IVDR emphasize => this indicated Annex refers to classification rules but not to the role of a manufacturer | Change to: Annex XIII | accept
3 | MedTech Europe | 177 ff (e.g.) | ed | References/citations could be moved to the footnotes unless absolutely necessary. This may improve reading flow and accessibility of the text (eg. Line 170) |  | accept
3 | MedTech Europe | 178-205 | Ed | Could be shortened - double information | Shorten the information | accept
3 | MedTech Europe | 186-205 | ed | This section could be rearranged and shortened to make it clearer in aim/purpose, definitions and methods. | Revise the section to make it simpler | accept
3 | MedTech Europe | 197-201 | ed | This paragraph mentions goal/aim and could be moved to the beginning of the section | Move this paragraph to the beginning of this section | accept
3 | MedTech Europe | 182 | Ed | Not entirely clear what is meant by “suitable registers” | Define “suitable registers” | accept
3 | MedTech Europe | 182 | Ed | Clarify wording. ‘does not need to but can’ suggests that it is optional to be proactive | The post market system should proactively collect and evaluate such clinical/performance information andthereby demonstrate to actively monitor information from post-market experience. | accept
3 | MedTech Europe | 184 | ed | Minor editorial change: delete the word “to” | Change to: “and evaluate such clinical/performance information and thereby demonstrate actively” | accept
PMS as an integral part o | MedTech Europe | 217 | Te | Not entirely clear what is meant by “similar products” | Define “similar products” | accept
PMS as an integral part o | MedTech Europe | 220 | ge | Regarding the usability issues: is the intention here to have a specific section for use errors? | Clarification in needed on the scope of “usability issues” | accept text revised
PMS as an integral part o | MedTech Europe | 221-222 | Te | The examples provided don’t usually have set black and white thresholds. There are triggers for further investigations and activities such as CAPAs and Escalations determine whether field actions or design changes are required | Remove eg. “starting a field safety corrective action, issue a design change etc.) | accept text revised
PMS as an integral part o | MedTech Europe | 212 | Ge | Identify the “amount” of data. Is it required to identify the amount of data? | Remove “amount of data” requirement. | accept
PMS as an integral part o | MedTech Europe | 211-231 | ed | The paragraph could be shorter, structured better and bullet point could be added | · e.g. point 1 and 2 overlap and are hardly distinguished, point 3 is afterthought. And so on… e.g. point 2 seems to use “performance” as non-defined general term and not opposed to “benefit” | accept text revised
PMS as an integral part o | MedTech Europe | 212-231, 253-260, 269 | ed | Line 212 – 231 AND Lines 253- 260 full stops inconsistency in bulleted sections Line 269 full stop missing | Text inconsistent. Edit full stops and remove inconsistencies. | accept
PMS as an integral part o | MedTech Europe | 228 | Ed | Is the word “publicly” missing for “utilization of available information” | The PMS plan shall address the collection and utilization of publicly available information and shall describe the systematic process to actively collect this information. | accept
4 | MedTech Europe | 237 | ed | Text not consistent with Article 83 MDR | Replace “requests” with “states” | rt 83 and 80 are 'shall' articles request is more accecable that state that the manufacture shall..
4 | MedTech Europe | Sections 4 to 5 (lines 234 – 486) | ed | The two sections contain a lot of repetitions. The 10 pages could be halved. I also recommend increasing the logical structure of both sections. This would avoid the repetition of information and mainly and foremost will increase the readability and comprehensibility. | 1. Recommendation: PMS system and components Main activities Interactions with other processes | accept doubling wil be adresses
4 | MedTech Europe | Sections 4 to 5 (lines 234 – 486) | ed | The two sections contain a lot of repetitions. The 10 pages could be halved. I also recommend increasing the logical structure of both sections. This would avoid the repetition of information and mainly and foremost will increase the readability and comprehensibility. | · Recommendation: PMS system and components Main activities Interactions with other processes | comment is a dublicate to line 237, can be taken out.
4 | MedTech Europe | Sections 4 to 5 (lines 234 – 486) | ed | The two sections contain a lot of repetitions. The 10 pages could be halved. I also recommend increasing the logical structure of both sections. This would avoid the repetition of information and mainly and foremost will increase the readability and comprehensibility. | Recommendation: PMS system and components Main activities Interactions with other processes | comment is a dublicate to line 237, can be taken out.
grafic 1 | MedTech Europe | 241 | ed | Figure lacks descriptor/legend |  | agreed, graphic will be taken out or replaced with a different graphical representation.
grafic 1 | MedTech Europe | 241 | ed | Figure lacks structure and relevance. Table could be more helpful with structure like “detect/contribute/update”. In current format hard to grasp and figure does not add information that could not be captured in text or table. | Figures should be used if content is too complex for text. e.g., complex flowchart of interactions of documents like SSCP, CER, PSUR etc. Here, much more relevant information could be introduced. Why is “if applicable” only added to SSCP? This seems inconsistent . It seems, the figure intends to summarize the following sections, but without structure or explanation. | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | MedTech Europe | 258 | ge | Would be better to align with 14971 standard | whether the collected data has any impact on the probability and the severity | accepted.
4a | MedTech Europe | 262 | ge | Are the indicators and threshold values aligned with the ones in trending as per Art. 88? | Please, clarify if this is aligned with Art. 88 | declined - will be explained in the Trend Reporting Guidance and a footnote in the PMS Guidance will refer to the Trend Reproting Guidance.
4a | MedTech Europe | 264 | ge | Is this MDR specific or is there an IVDR ref missing? | Incorporate IVDR: "(as referred to in Section 3 of Annex I MDR and IVDR)?” | accept
4a | MedTech Europe | 243-269 | ed | Aim and content of the section are vague. Very general information and lack of clear guidance (i.e. intent of the document) Clearly focus on the aim/ title of the section (same applies for the following sections) | Aim: Update Benefit-risk- assessment/determination Content: clarify methods and guidance beyond generalities | noted
4a | MedTech Europe | 248 | ed | Remove a duplicating word | Change to: “acceptable. Furthermore, the data should be utilized in the improvement of the risk.” | accept
4a | MedTech Europe | 257 | ed | Minor editorial change: delete “in order” and replace with “to” | Change to: “market surveillance activities and the appraisal of the data to determine” | accept
4a | MedTech Europe | 263 | ed | Remove article ‘the’: …PMS Plan and be the linked | Change to: “…PMS Plan and be linked” | accept
4a | MedTech Europe | 263 | ed | Grammar correction needed: …and be the linked… | Delete ‘the’ | accept
4a | MedTech Europe | 266-269 | Ed | Rephrasing the whole paragraph suggested | Recommend rephrasing: ”The post-market surveillance system should include processes to identify, document, and evaluate changes to the benefit-risk profile of the medical device for acceptability. The post market surveillance system must include processes to initiate appropriate measures, including preventive and/or corrective action as required. “ | partially accepted - will not be inserted in this section but TF will review where this verbiage could add value in the document.
4c | MedTech Europe | 285 | Te | Wrong reference: ‘Art. 78.3 (b) IVDR’ | Change to: Art. 78.3 (c) IVDR | reject
4c | MedTech Europe | 296 | Ge | Include benefit-risk profile as well? | Change to: “ensuring the continued acceptability of identified risks, benefit risk profile, and” | accept
4c | MedTech Europe | 287-290 | ed | Text in this paragraph (lines 287 – 290) difficult to read | Perhaps. This could be separated into 2 sentences | accept, has been rephrased.
4c | MedTech Europe | 297 | ed | Verbiage should align directly to EU MDR: “detecting emerging risks on the basis of factual evidence” | Change to: “identifying and analysing emergent risks on the basis of factual evidence” | accept
4c | MedTech Europe | 302 | Ed | IVDR PMPF performance follow-up should be used instead of clinical follow-up. | Change post market clinical follow-up to post market performance follow-up for IVDR. | accept
4c | MedTech Europe | 308, Footer page 8 | ed | Footer on page 8 guidance Q&A on trendreport | Change “trendreport” to “Trend Reporting” | accept
4c | MedTech Europe | 311 | ed | Minor editorial change: ‘I’ is missing on “Identifying” | Change to : “Identifying” | accept
4c | MedTech Europe | 319 | ed | Line 319 Annex III, 1. (b) Indent 10 missing information | Annex III, 1.1 (b) Indent 10 | Reject, see latest version of the MDR.
4c | MedTech Europe | 330/331 | ed | Missing a word | Change to: “report / performance evaluation report and the technical documentation. The conclusions” | Reject - sentence has been restrcutured.
4d | MedTech Europe | 348-349 | te | Change of occurrence that leads to risk level increase to high or all changes? | Clarify the scope | Accept - sentence rephraed.
4d | MedTech Europe | 351-352 | te | Do we need to include high level summary of PMS data in SSCP? | Clarify the scope | Partially accept - will be rephared. Maybe add link to SSCP guidance.
4d | MedTech Europe | 350 | te | Include Performance Evaluation Report in addition to CER to account for IVDR |  | accept
4d | MedTech Europe | 333-334 | ed | Line 333-334 (referred to in Section 3 of Annex I). | Terminology: is it correct to use the word “Section”? | Reject, "Section" is correct. Article 10, 9(e) MDR.
4d | MedTech Europe | 350-351 | ed | Line 350-351 Periodic Summary Update report – | Such a report does not exist. There is a PSR (Periodic Summary Report) and a PSURs (Periodic Safety Update Report) Not clear what is meant here? | accept
4d | MedTech Europe | 351 | ed | Line 351 – full stop missing after report | Add full stop after “report” | accept
4e | MedTech Europe | 358 | ed | Minor wording change | Change to: “…entire lifetime, in addition, to drawing…” | not relevant -text is changed
4f | MedTech Europe | 372 – 374 | ge | Clarify if this ‘consideration’ needs to be documented somewhere. Current wording makes this optional |  | accept - replace with "use"
4f | MedTech Europe | 369.372 | ed | Minor editorial changes | Change to: “requirements for receiving, reviewing, and evaluating information, including feedback and”Change to: “parties. The information should be considered to identify any options to improve” | accept
4f | MedTech Europe | 372 | ed | Minor editorial changes: article ‘are’ is missing in “…if there any” | Change to: “…if there are any” | accept
4f | MedTech Europe | 374 | Ed | Refer to “usability of the device” instead of “use of the device” | Change to: “The information should be considered to identify if there any options to improve the safety or performance of the device and the feedback could be used to improve the usability of the medical device.” | accept
4g | MedTech Europe | 376 | ge | Clarification needed if this only pertains to other devices held by the manufacturer |  | accept paragraphs are rephrased
4g | MedTech Europe | 376 - 385 | ge | The whole paragraph seems to go beyond the MDR text: The following section is introducing requirements which are not present in the EU MDR. The MDR text is not mandating the manufacturer to demonstrate actions taken based on PMS information from one device to another. In addition, Art 83.3 (g) specifies “when relevant”: ((g) when relevant, to contribute to the post-market surveillance of other devices) which enables the manufacturer to apply relevancy criteria to the use of the data. |  | Accept to re-introduce the relevancy in the title - the paragraph should be rephrased - the regs do not speak about "similar" devices; only "other" devices. Paragraph needs to be rephrased - Sietske to propose wording
4h | MedTech Europe | 390 | te | As per Article 88(1), additional criteria is necessary for determination of analysis requirement. | Before “should”, insert “that could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits.” | rephrased
5 | MedTech Europe | 400 | ed | Figure 1 missing in the document | Add Figure 1 | accept
5 | MedTech Europe | 403, 404 | ed | Wrong reference - 6.1.1 and 6.1.2 | Change to: 5.1.1 and 5.1.2 respectively | accept
5.1.1 | MedTech Europe | 413 | te | There is no data related to manufacturing process being evaluated as part of PMS in this section | Include manufacturing process in the scope | reject, manufacturer proces deviations often results in CAPA's or FSCA and those are already in the tabel
5.1.1 | MedTech Europe | 414 | te | As per MDR article 88, “Manufacturer shall report (…) any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side- effects that could have a significant impact on the benefit-risk analysis”. Therefore, trend reporting is applicable to non-serious incidents. In row 2 “Information concerning serious incidents, including information from PSURs, and field safety corrective actions” of table of paragraph 5.1.1 it is stated that “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR This information should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR. ”. | Delete the sentence regarding trend reporting for serious incidents | accept
5.1.1 | MedTech Europe | 414 | te | Information concerning serious incidents, including information from PSURs, and field safety corrective actions: How to utilize it: “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR” But according Article 88 MDR: "any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side- effects that could have a significant impact on the benefit-risk analysis”. So, Trend reporting should only include “serious expected undesirable side- effects” but not “serious incidents” in general. | Clarification would be appreciated | accept, will be reworded.
5.1.1 | MedTech Europe | 414 | te | “A summary of any new or emerging serious incidents along with any field safety corrective actions issued should result in an update of the periodic safety update report (PSUR) in line with the PSUR guidance.” First, I don’t understand the sentence. If you take out “along with any field safety corrective actions issued”, the sentence doesn’t make sense. A summary can be provided in the report, but how can it result in an update? Second, there is nothing in the PSUR guidance that indicates that the PSUR needs to be updated in these cases. | A summary of any new or emerging serious incidents along with any field safety corrective actions issued should be provided in the periodic safety update report (PSUR) in line with the PSUR guidance. | accept
5.1.1 | MedTech Europe | 414 Page 11 | te | Manufacturers are expected to - collect feedback and complaints related – pg 11 | Remove | accept
5.1.1 | MedTech Europe | 414, column 2, row 2 | te | “ For example, when the manufacturer plans a survey, there should be a plan for the design and execution of survey, with clearly identified objectives and evidence that it addresses appropriate questions relating to the post-market experience of the medical device. “ | Too detailed and specific. Plan how to conduct survey is not part of PMS Plan | accept
5.1.1 | MedTech Europe | 414, column 2, row 3 | te | “ There is an expectation that manufacturers have robust procedures in place to collect and track information related to serious incidents and complaints. It is not entirely clear what is meant by Complaints in this context | This ANNEX III requirement only speaks of Serious Incidents, FSCA and PSURs. PSUR is unusual in this context, as it in turn summarizes the data on serious incidents and FSCAs. | accept - moved to "Records referring to non-serious incidents and data on any undesirable side-effects"
5.1.1 | MedTech Europe | 414, column 2, row 7 | te | “ Manufacturers are expected to be proactive in gathering feedback and complaints” Not entirely clear what is meant with this statement? | If complaints are defined according to ISO 13485, should the manufacturer actively ask about these complaints??? Feedback yes, but complaints are a result | accept
5.1.1 | MedTech Europe | 414, column 2, row 8 | ge | Expectation to screen CA websites outside EEA on “similar devices” – imagine the complexity of doing this in every language in the world. | Suggested to remove “including those outside of the EEA” | accept
5.1.1 | MedTech Europe | 414, column 3, row 3 | te | “ Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. “ The statement is not entirely accurate. | Incidents should be used instead of serious incidents | accept
5.1.1 | MedTech Europe | 414, column 3, row 3 | te | “ A summary of any new or emerging serious incidents along with any field safety corrective actions issued should result in an update of the periodic safety update report (PSUR) in line with the PSUR guidance.” The statement is not entirely accurate. | The PSUR Guidance does not speak of event- triggered updates. On the contrary, the manufacturer should adhere to the 12-month rhythm and have to justify when you make an interim update. | accept
5.1.1 | MedTech Europe | 414, column 3, row 5 | Te | “The outputs of trend reporting should be used to notify Competent Authorities in relation to articles 88 and 89 of the MDR and articles 83 and 84 of the IVDR. The manufacturer should consider whether field safety corrective actions are required to be issued including field safety notices” Article 89 is not related to trend reporting | Article 89 should be removed | accept -TF on Trend Reporting will be consulted to ensure alignment
5.1.1 | MedTech Europe | 414, Page 12 | te | “How to utilize it” row titled “Information from trend reporting” should be consistent with Article 88(1). | Replace “The outputs of trend reporting should be used to notify Competent Authorities” With Trend reporting shall be submitted to Competent Authorities by means of the electronic system referred to in Article 92 MDR” | accept -TF on Trend Reporting will be consulted to ensure alignment
5.1.1 | MedTech Europe | 414, Page 11 | ge | “How to utilize it” row titled “Information concerning serious incidents, including information from PSURs, and field safety corrective” The text refers to Article 88 MDR but Article 88 focuses on any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side-effects that could have a significant impact on the benefit-risk. The guidance is confusing trending, trend reporting, reporting on all PMS data. Can we refrain from using the word 'trends' when we are not talking about trending specifically related to trend reporting or make it clear in that section of the guidance? |  | Noted - TF will check the entire document before finalisation for the useage of wording "trend"
5.1.1 | MedTech Europe | 414, Page 11 | ge | Page 11 The post-market surveillance plan should address how feedback and complaints are received and processed by the manufacturer, as well as the methods and tools that that are used to investigate complaints and to analyse market-related experience collected in the field. Manufacturers are expected to - collect feedback and complaints related to the medical device. The manufacturer may therefore consider using research methods such as surveys to collect feedback from users, distributors, importers and other relevant economic operators. For example, when the manufacturer plans a survey, there should be a plan for the design and execution of survey, with clearly identified objectives and evidence that it addresses appropriate questions relating to the post-market experience of the medical device. | Is the guidance trying to mandate this information to be captured in the PMS plan? Should this information not be in a procedure rather than a plan? | accept - same comment as in 506
5.1.1 | MedTech Europe | 414, Page 12 | ge | Information on trend reporting pg 12 See further the Q&A on trend reporting | Is the MDCG Q&A on Trend Reporting being supplied? Not possible to comment and evaluate without this document being available. | Text will be aligned with available MDCG Guidance on Trend Reporting.
5.1.1 | MedTech Europe | 414, Records referring to non-serious incidents and data on any undesirable side-effects | Ge | Should the reference to article 88 for trend reporting be added? Also include the applicable IVDR reference | Change to: “In relation to non-serious incidents and undesirable side-effects, manufacturers should ensure that the exchange of information between users of the device (or other economic operators) and the manufacturer is facilitated by processes that allow for easy and timely trend reporting as required per Article 88 (MDR) and Article 83 (IVDR).” | partially accepted - agree to update text in line to include requirements from IVDR.
5.1.1 | MedTech Europe | 414/point 5.1.1. Table row 2: Information concerning serious incidents, including information from PSURs, and field safety corrective actions | ge | Article 88 refers to trend reporting of incidents that are not serious incidents or that are expected undesirable side- effects, but current text mentions that serious incidents should be used for trend reporting. Current text: How to utilize it: Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. This information should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR. Per Article 88: Manufacturers shall report, by means of the electronic system referred to in Article 92, any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side- effects that could have a significant impact on the benefit-risk analysis | Remove the first sentence of “how to utilize it” (highlighted in blue) | accept
5.1.1 | MedTech Europe | 414 | ed | Table of 5.1.1 shows duplicated information on page 11 and 13: ‘Information, including feedback and complaints, provided by users, distributors and importers;’ | Remove duplicated information | accept
5.1.1 | MedTech Europe | 438 | ed | Section 5.1.1 does not include all inputs mentioned in figure 2 | Include the missing points | it is good to align a table with the mentioned actions in 4.1.2
5.1.1 | MedTech Europe | 414 ff. | ed | Table does not seem to be fully accessible and the text is slightly too long. Perhaps, it would be good to reduced by 70-80%. Also, the aim of the table is not entirely clear. Distinction between “collect” and “utilize” is unclear. is the content could be more focused. |  | accept
5.1.1 | MedTech Europe | 414, Table 5.1.1 Rows 1 and 6 | ed | Section on information including feedback and complaints, provided by users, distributors and importers appears twice in this table. | Please remove the 2nd item if duplicate or update the descriptions to distinguish the required information | accept
5.1.1 | MedTech Europe | 414, Page 12 | ed | Under column “How to utilize it” row titled “ Records referring to non-serious incidents and data on any undesirable side-effects ” should be consistent with 348- 349. | Replace “emerging” with “increased” | accept
5.1.1 | MedTech Europe | 414, Page 12 | ed | Under column “How to collect it” row titled “Information from trend reporting” should be consistent with Article 88(1). | Remove “. This includes those” | Trending TF will be consulted.
5.1.1 | MedTech Europe | 414, Page 12 | ed | Under column “How to utilize it” row titled “Relevant specialist or technical literature, databases and/or registers” should be consistent with 348-349. | Replace 2X “emerging” with “increased” | accept
5.1.1 | MedTech Europe | 414, Page 12 | ed | Information on trend reporting pg 12:GSPR abbreviation is not define before hand | Suggestion to add GSPR abbreviation when General Safety and Performance Requirements is first mentioned in the text | accept
5.1.1 | MedTech Europe | 414, Page 13 | ed | Under column “How to utilize it” row titled “Publicly available information about similar medical devices” should be consistent with 348-349. | Replace “emerging” with “new or increased” | accept
5.1.1 | MedTech Europe | 414, Page 13 | ed | Page 13:“The post-market surveillance plan should explain how feedback and complaints are received and processed by the manufacturer, as well as the methods that are used to actively gather feedback not only from users of the device but also economic operators such as distributors and importers.” | The language used seems to be quite complex and not entirely clear. Suggestion to put it in a simpler way: Manufacturers should explain the process for processing feedback. | Setence was adjusted
5.1.2 | MedTech Europe | 419, Page 15 line 2 column 2 | ge | “Procedures cover the internal procedures as well the procedures to communicate with users, economic operators, notified body and competent authorities.” This is redundant with page 14, last line | Procedures cover the internal procedures. The procedure to communicate with users, economic operators, notified body and competent authorities are already described in the line “methods and protocols to communicate effectively with competent authorities, notified bodies, economic operators and users;” | accept
5.1.2 | MedTech Europe | 419, effective tools to trace and identify devices for which corrective actions might be necessary; and | Te | We are not able to identify where the requirement to identify product status is applicable from MDR with respect to post market surveillance. | Remove language for identifying product status. | accept
5.1.2 | MedTech Europe | 419, Page 14 | ge | Section row “effective and appropriate methods and processes to assess the collected data” talks about effective and appropriate methods and processes to ASSESS collected data, but then the guidance only really highlights the appropriateness of the data in relation to the classification (ie specific PMCF for a class III) but nothing to really indicate what are considered acceptable methods for assessment. The guidance highlights the importance of ANALYSING the data but then doesn't really add any clarification on how to utilize the data / ways to analyse / thresholds and indicators and their importance when feeding it into the risk management documentation. This appears to be a miss in the guidance in terms of highlighting the importance of and the relationship with risk management |  | accept
5.1.2 | MedTech Europe | 419, Page 14 | ge | Section row “suitable indicators and threshold values1 that shall be used in the continuous reassessment of the benefit- risk analysis and of the risk management as referred to in Section 3 of Annex I this text doesn't really say anything to further clarify the requirement. Please provide examples or scenarios to establish what is considered as appropriate thresholds and indicators for analysis of collected data.”: | Please provide examples or scenarios to establish what is considered as appropriate thresholds and indicators for analysis of collected data. | reject - this can not be achieved by the guidance but should be adressed by the manufacturers procedures for risk management.
5.1.2 | MedTech Europe | 419, reference to procedures to fulfil the manufacturer s obligations laid down in Articles83, 84 and 86 MDR and article 78, 79 and 81 IVDR; | Ge | Should we add a reference to MDCG 2022-21 for PSUR content. | Change to: “The manufacturer should clearly identify the relevant procedures for post-market surveillance activities as described in Article 78, 79 and 81, including the PMS plan, the data to be collected and recorded, its analysis and the requirements for the PMS report / PSUR. Reference MDCG 2022-21 for PMS report and PSUR guidance.” | reject - text was modified.
5.1.2 | MedTech Europe | 419, systematic procedures to identify and initiate appropriate measures including corrective actions; | Ge | Since there is guidance on Corrective and Preventive Actions, should we include guidance on FSCA as well? | The procedures for actions to be taken including corrective actions, including Field Safety Corrective Actions, should be identified. | accept - text has been modified.
5.1.2 | MedTech Europe | 414, row 4 | te | “Records referring to non-serious incidents and data on any undesirable side-effects”: Does the intent is to trend separately non serious and serious event? | Clarify the scope | noted - Text will reference the TR Guidance
5.1.2 | MedTech Europe | 419 | te | Row labelled ”— systematic procedures to identify and initiate appropriate measures including corrective actions;” explanation appears to address a different requirement. | Add explanation to address requirement | tabel is revised
5.1.2 | MedTech Europe | 419 | te | Row labelled ”— a PMCF plan as referred to in Part B of Annex XIV, or a justification as to why a PMCF is not applicable. (MDR)” explanation second bullet should be consistent with 348-349. | Replace “any increased risk” with “new or increased risk” | tabel is revised
5.1.2 | MedTech Europe | 419 | ed | Row labelled ”— effective tools to trace and identify devices for which corrective actions might be necessary; and” explanation typo. | Replace “identifying” with “identification” | accept
5.1.2 | MedTech Europe | 419, Page 15 second column | ed | Confirm the safety and performance profile of the device where the initial approval was based on equivalence If the manufacturer determines that no PMCF/PMPF activities are required, the manufacturer should provide a justification within the post-market surveillance plan. I think the first sentence should be a bullet point like the sentences before. There should be a dot before “If the manufacturer”, and I think this sentence should be in a different paragraph. | Change to: Confirm the safety and performance profile of the device where the initial approval was based on equivalence. If the manufacturer determines that no PMCF/PMPF activities are required, the manufacturer should provide a justification within the post-market surveillance plan. | accept
5.1.2 | MedTech Europe | 419, Page 14 | ed | Pg 14 . refer to the trend Q&A – reference 8 which is not added | But reference 8 should be added | accept
5.1.2 | MedTech Europe | 419, Page 15 | ed | Pg 15 Articles83, — reference to procedures to fulfil the manufacturers obligations laid down in Articles83, 84 and 86 MDR and article 78, 79 and 81 IVDR; | Are these the correct articles? | Accepted - Art 83-86 and art 78-81
5.1.2 | MedTech Europe | 419/table row 10, column 2 | ed | Word ‘of’ needs to be removed in last sentence: “well and procedures for identifying of the relevant economic operators” | Remove ‘of’ | accept
5.1.2 | MedTech Europe | 419/table row 6, column 2 | ed | Last sentence not clear: “More about trend obligations you can consult the MDCG Q&A trend. refer to the trend Q&A” | Change to “For more about trend obligations refer to the trend Q&A.” | accept
5.2 | MedTech Europe | 422 | Ge | This section is duplicate to the content in the beginning of the guidance. Should this be removed to simplify the document? | Remove section 5.2 | reject chapter 4 is explaining all steps, a bit of duplication from the introdudctional cahpter are possible.
5.2 | MedTech Europe | 437 | ge | Fig. 2: the relevance, content, and design is not entirely clear. What is the aim? What do the larger overlapping circles mean? Is it supposed to be some chart? Circles with content does not seem to be clear enough. The structure could be improved. Also, the figure is partially not based on regulations. It is not clear where some of the terms are coming from (e.g. “focus groups”)? These are not defined and standardized. |  | accept, figure is deleted
5.2 | MedTech Europe | 482 | ge | MDGC 2022-21 does not specify that the next period of PMS starts with the review of the plan. If this isn’t included in 2022- 21, is it appropriate to include here? | Remove the following sentences: “the PMS report or the PSUR8.” “to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes.” | accept
5.2 | MedTech Europe | 435 | ed | The statement appears to be in regard to devices with higher classification. | Replace “higher-risk” with “higher benefit- risk profile” | noted
4d | MedTech Europe | 348 | Te | Define ‘new information’ – the pseudo definition given in parentheses includes ‘safety/performance information obtained from PMCF/PMPF’, but that data may not impact the B/R profile. So, an SSCP might not have the latest numbers from PMCF/PMS because there was no change to the B/R, but lines 349 - 351 says the SSCP needs to be aligned to the CER and PSUR. Mixed message |  | rejected, simplefied the sentence.not for PMS guidance, but for the SS(C)P guidance
5.3 | MedTech Europe | 456 | te | Line 456 plan (Article 85). | Should this be Article 84? | delete the reference?
5.3 | MedTech Europe | 440 | ed | Whole section is redundant information and general advice. | Should be rewritten or removed. | perhaps make clear that PMS system is a plan do check act? Then 5.2, 5.3 and 5.4 are mor logic?
5.4 | MedTech Europe | 479 | te | Not clear what kind of ‘Summary report’ is meant. | Clarification needed, what kind of ‘Summary report’ is meant. | accept
5.4 | MedTech Europe | 482-483 | ge | “to start the next period of PMS, a review of the PMS plan is necessary”: such review is not mandatory as per IVDR | The manufacturer shall update the PMS plan as needed. | accept
5.4 | MedTech Europe | Line 482 – 485 | ge | “to start the next period of PMS, a review of the PMS plan is necessary.” There is no PMS period as it is a continuous process. And this is not a requirement from the regulation | Delete the paragraph | accept
5.4 | MedTech Europe | 464 | ed | Whole section is redundant information and general advice. | Should rewritten or removed. | perhaps make clear that PMS system is a plan do check act? Then 5.2, 5.3 and 5.4 are mor logic?
5.4 | MedTech Europe | 469 | ed | Minor editorial changes | Change to: “-Articles 83 (MDR) and Article 78 (IVDR) make it clear that it is mandatory for all…” | accept
5.4 | MedTech Europe | 476 | ed | Something seem to be missing in a sentence: “So besides planning the PMS, and follow the procedures for collecting information.” |  | accept
5.4 | MedTech Europe | 476 | Ed | Sentence seems to be incomplete |  | accept
5.4 | MedTech Europe | 476 | ed | Confusing or partial sentence | Rewrite sentence | accept
5.4 | MedTech Europe | 479 | ed | Minor editorial change: I think the dot and the “8” should be parenthesis. | Change to: “Conclusions need to be drawn up in a summary report (the PMS report or the PSUR).” | accept
5.4 | MedTech Europe | 479.48 | ed | There are excessive words in a sentence: “Conclusions need to be drawn up in a summary report. the PMS report or the PSUR8.” | Remove ‘a summary report.’ and ‘8’ | accept
5.4 | MedTech Europe | 482 | ed | Letters are not capitalized at the beginning of the sentences: “to start the next period of PMS, a review of the PMS plan is necessary. this review is to” | Capitalize ‘To’ and ’This’ at the beginning of the sentences | accept
5.4 | MedTech Europe | 473 - 485 | ed | Complete sentences seems to be missing, in line 480 “PSUR8” | Please check for complete sentences and check the foot note. May be 9 is correct. | accept
5.4 | MedTech Europe | 476 - 480 | ed | Minor editorial changes | Change to: “-So besides planning the PMS, and following the procedures for collecting information, the data needs to be analysed and assessed. | accept
5.4 | MedTech Europe | 476-480 | ed | Paragraph Line 476 – 480 is not clearly written | This text doesn’t read correctly. Please rewrite in a clear and simple manner to make it easier to understand. | accept
5.4 | MedTech Europe | 482-485 | ed | Line 482 -485 to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes. | Confusing or broken sentences/section. This text doesn’t read correctly. Please rewrite in a clear and simple manner that it is understandable. | accept
6 | MedTech Europe | 505 | ed | Minor editorial change in Line 505: “.the” | Capital letter ‘T’ should be used | accept
6 | MedTech Europe | 489-491 | ed | The paragraph (Line 489-491) is not clearly written | It is difficult to read. Propose to re-write in a clear and simple way. | accept
annex I | MedTech Europe | 529 PMS Report - MDR Article 85 / IVDR Article 80 | Te | Provide an example for “when does a PMS report need to be updated”. What are conditions that may be considered necessary. |  | partially accepted. Text modified.
annex I | MedTech Europe | 529, Annex I Table, PSUR section | ge | The timing included for updating PSUR documents for Implantable devices of at least annually does not match the PSUR guidance (MDCG 2022-21) for Class IIA products of at least every 2 years. | Updated Class IIA and IIB device sections to include implantable devices and have Class III devices as an independent section. | accepted
annex I | MedTech Europe | 529, Table, PSUR - MDR Article 86 / IVDR Article 81 | te | The table reads: “IVDR: applicable to class C and D devices”. It should be clarified that PSUR is not applicable for IVD legacy devices. | Add clarification (perhaps via a footnote): “PSUR is applicable for all IVDR devices, except devices which comply with IVDR Article 110.3 (IVD “legacy devices”).” | partially accepted. Table revised and pointing to MDCG 2022-21 now instead.
annex I | MedTech Europe | 529, PMS Plan - MDR Article 84 & Annex III / IVDR 79 & Annex III | Ge | Under the obligations column, would be good to include a bullet for complaint handling/reporting specifically. | Add the following bullet: “The methods and tools to investigate complaints and analyse market-related experience collected in the field;” | accepted
annex I | MedTech Europe | 529 | ge | Table section : "PMS Report - MDR Article 85 / IVDR Article 80", "Obligations" column " What information must a PMS Report contain at a minimum?” The NB Group position paper indicates that information such as sales data etc should also be included in the PMS report to give context to the report. Therefore, in terms of content there’s no difference between a PMS Report vs. PSUR. This is the reference to the text for the Team-NB Position Paper for IVDs in the Post Market Surveillance Section: "The manufacturer should also provide the following post market surveillance data for the last 5 years: Market History. Worldwide and EU sales volumes. Complaint data and trend analysis. Vigilance data and trend analysis (include details of any adverse incidents, recalls or FSCA) Publicly available information about similar medical devices. Modifications made and/or corrective actions taken following the incidents reported and revisions made to the risk management file." | · a summary of the results and the conclusions of the analyses/assessment of the post-market surveillance data defined in the PMS plan. A rationale and description of any preventive or corrective actions taken during the period covered by the report. conclusions on the safety and the performance of the device. data on sales volumes and estimates of user population and frequency of use. a rationale and description of any preventive or corrective actions taken during the period covered by the report. | partially accepted. Reference to MDCG 2022-21 added.
annex I | MedTech Europe | 529, 529, Page 21 | ed | Minor editorial change in Pg 21: “Implantable and Class III devices:• Submitted by means of EUDAMED to notify body.” | Change to: “notified body.” | accept
annex I | MedTech Europe | 529, Page 18 | ed | Minor editorial change in Pg 18 table: “78Articles” | Change to: “78 Articles” | accept
annex I | MedTech Europe | 529, Page 21 | ed | Pg 21: “PSUR and NB evaluation available to CA in EUDAMED” | Change to: “PSUR document and Notified Body evaluation available to competent authorities in EUDAMED.” Rationale: Avoid usage of abbreviations and use clear verbiag. | accept
ge | CA6 |  | Ge | General comment: Lacking guidance on fundamental concepts of adapting a PMS system for every product, e. g.: • Timing of activities; continuous vs planned for a certain timepoint • Timing: How could the PMS system change over time for a particular device; when it is newly developed vs when it has been on the marked fof many years? • Timing: How and where to describe timing for PMS activities, e. g. planned activities in the PMS plan vs the first version and later updates of the corresponding PMS report/PSUR? Should there be a pre-set cycle? If so, how to set this timing? • Grouping of similar devices into one PMS plan • Adaptation: What would be “suitable”, “effective” and “appropriate”, e. g. methods, thresholds and indicators, … • Adaptation: Which aspects of a PMS system should be applied differently for different devices? Which aspects can vary and what would be proportionate to the risk class? and appropriate for the type of device?
intro | CA6 | 58 – 77 | te | Suggest adding text somewhere in the introduction which clarifies that the term “medical device” in this document also includes IVD medical devices. |  | Accept - removed from definitions and added to introduction
intro | CA6 | 76 | Te | Is the scope of PMS really limited to GSPR compliance alone?? Suggested rewording to broaden the scope: | “… continues to comply with all applicable requirements in MDR/IVDR, which should be reflected in the technical documentation for each device”. | accept
1 | CA6 | 80 | ge | Is this really the purpose?? Isn’t it rather to guide manufacturers in developing, applying and maintaining their PMS system? A PMS system may be set up in a number of ways, and we should not limit this flexibility, but rather provide guidance on how to interpret the requirements for a PMS system. |  | reject we are not saying how or limiting flexibility; some elements are mandated by the regs (e.g. PSUR…) -
3 | CA6 | 137 | Te | We suggest to add section(s) to explain that a manufacturer can have one PMS procedure but several PMS plans, such that each device is covered by a PMS plan, which can also cover a group of similar devices. This is a central principle that manufacturers don’t always understand and should be explained in a PMS guidance document. |  | reject - the question of grouping is valid but the guidance is not the place to discuss; maybe a Q&A will be required
PMS as an integral part o | CA6 | 223 | Te | We would find it useful to include that the manufacturer should have procedures to gain access to the device to obtain the necessary information on the quality, performance and safety of a device and to draw necessary conclusions… |  | accept text revised
4 | CA6 | 234 | te | (4. The interactions of the PMS processes with other key processes of the QMA) The section below does not describe or provide guidance on process interactions. Rephrase heading or content to match. Merely repeating regulation text does not provide sufficient guidance. |  | reject is an intruduction for this chaper
grafic 1 | CA6 | Figure 1, row 240 | ge | The figure only repeats information that is already in the MDR/IVDR and does not add anything new on how the processes should interconnect, or even which processes that are meant. | More interpretations would increase the usefulness of the figure. | agreed, graphic will be taken out or replaced with a different graphical representation.
4c | CA6 | 318-319 | Ge | This is something that should be expanded and explained in a guidance such as this one! What differs these two plans from each other? How do they relate to individual or groups of products? How do/could they change over time? |  | agreed. Text will be amended to provide further clarity. Will explain more about what should be in the PMCF Plan and will insert reference to MDCG 2020-7 (PMCF Plan Template).
4e | CA6 | 354 | Ge | The text below does not provide any guidance or examples on how such needs could or should be identified. Examples of each step in this section would be very beneficial. |  | reject The "how" is not part of the guidance; this will be done per analysis conclusion and manufacturer processes;
4f | CA6 | 368-374 | Ge | The text should but does not contain any information relating to usability (IEC 62366). |  | Partially accepted: mention of standard in guidance is not supported - text changed
4g | CA6 | 382 | Te | Design characteristics: This central concept needs clarification in this guidance, or reference to somewhere else if it has already been explained elsewhere. |  | reject- define design characteristic not in scope of the guidance
5 | CA6 | 399 | ge | “5. the main activities within a PMS system” It is important that this section does not only apply to one product, one PMS plan, but that it also applies to more complex organisations and product portfolios |  | reject -the PMS of the manufcaturer should describe how to handle complex situations
5.1 | CA6 | 406 | Ge | 5.1. Planning This section is lacking information on PMS plans and planning for MFs with a larger product portfolio. | Include text to answer questions such as: How can devices be grouped? And how to multiple PMS plans relate to one PMS system? Can a MF add and/or delete devices to a PMS plan? How and when? And how does this relate to the requirements for PMS reports and PSURs? | seek legal advice from COM Laywer to understand PMS "for each device"
5.1 | CA6 | 410-412 | te | For purpose of introducing more explanation on how the concept of PMS plans can support effective PMS for each device, please consider rephrasing or adding a subordinate clause. | The post market surveillance plan for each device or group of devices, should be adapted based on the conditions to achieve effective post market surveillance for each device (group of devices). | reject, adressed in paragraph above
5.1.2 | CA6 | Table 1 415 | Ge | Lacking the very important aspect on how a manufacturer should adapt the PMS system and plan to be “proportionate to the risk class and appropriate for the type of device”. All guidance provided below is for one level only. |  | agree , adressed in ch 1
5.1.2 | CA6 | Table 1, 2nd row | Ge | Column “How to utilize it”, second row “information concerning serious incidents, including information from PSURs, and field safety corrective actions”: There is a need for interpretation of the regulations – do we mean that this section concerns one device, the manufacturers devices or all similar devices on the market irrespective of manufacturer? This section seems to limit the scope of information collection to only apply to the device in questions. The wording in the regulations does not make this limitation. Do we want to limit the scope to only cover each device in itself, and not to look for and act upon information, e. g. incidents and FSCAs, that is readily available for similar devices on the market? |  | This specific point is device specific. Information regarding other and similar devices is covered under "Publicly available information about similar devices".
5.1.2 | CA6 | Table 1, 2nd row | Te | “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR.” Incidents can be objects for trend reporting (art 88/83), but not serious incidents. | Information relating to incidents should be used for… | accept
5.1.2 | CA6 | Table 1, 3rd row | Ge | Lacking guidance on which information should be collected as non-serious incidents and undesirable side effects and any expectations on such procedures. |  | accept. Text modified. "There is an expectation that manufacturers have robust procedures to collect complaints and feedback, and evaluate if these complaints and customer feedback fulfil the criteria of an incident under Article 2(64) MDR/ Article 2(67) IVDR. Ref MDCG 2023-3"
5.1.2 | CA6 | Table 1, 7th row | Ge | “The manufacturer may wish to use several methods to collect this data, which could include types of data from the following (not exhaustive) list:” This way of providing guidance, in the form of suggestions, should be applied throughout this whole document. It is very helpful. |  | Noted - It is the intention of the TF to provide examples in a future update but not as part of the initial release.
5.1.2 | CA6 | Table 2, 2nd row | Ge | “For example, a class III implant is likely to require specific PMCF activities and the methods of analyzing the associated data should be appropriate for its risk classification.” Further guidance is needed. This is a very central concept that should be further extended in this guidance |  | partially accepted. TF will insert a reference to the relevant guidacne documents on PMCF ( MDCG 2020-8 -Guidance on PMCF evaluation report template MDCG 2020-7 Guidance on PMCF plan template
5.1.2 | CA6 | Table 2, 3rd row | Ge | “ The data obtained from these activities should also be used to set thresholds and indicators that will allow for the continuous reassessment of the safety and performance of the device…” Further guidance is needed. This is a very central concept that should be significantly extended in this guidance. |  | reject.
5.1.2 | CA6 | Table 2, 5th row | ge | “The post-market surveillance plan should describe the manufacturer’s processes and methods for investigating complaints and how to collect and analyse data related to market-related experience collected in the field” This text only elaborates on some of the requirements. No guidance is given on trending as such, statistical significance, frequency or severity scoring, or a suitable observational period or how this may or may not change over time. |  | rejected. Trend Reporting is a different requirement in Annex III 1B and will be discussed in the guidacne.
5.2 | CA6 | 422 | Te | Not clear why there are two headings. Also the text below does not provide further guidance related to any of these headings, neither how to assess not use data collected and analysed through PMS. |  | accept
figure 437 | CA6 | Figure 2 438 | Te | In order for this figure to fit with the text above, additional text must be added to create a logical bridge, e. g. explaining how these would be pro-active, compared to other sources which are re-active, e g. as in ISO 20416, section 5.5.2. |  | accept, figure is deleted
5.3 | CA6 | 466 | Te | “in particular be used for input and to update processes in the system” Don’t agree to this interpretation. The data collected through PMS shall be used to update the technical documentation for applicable product(s), through the processes in the QMS, e. g. risk management, clinical/performance evaluation, design/change control, etc. The aim is not to update the processes as such. |  | reject its the other way around: update the processes and The technical documentation shall be updated accordingly
5.3 | CA6 | 440 | Te | Not clear, why two headings? Also the text below does not provide further guidance related to any of these headings, neither how to assess/analyze or use data collected and analysed through PMS. |  | accept
annex I | CA6 | Annex I | Ge | General comment 1: The wording in this column is too concise and simplistic. It is easily misunderstood. Although a brief overview is appealing, the table should be carefully examined and reworded in order to ensure clarity. What are the obligations? |  | reject. This is beyond the scope of the docuemnt.
annex I | CA6 | Annex I | Ge | General comment 2: It would be more helpful if guidance could be provided for what would be considered as “suitable”, “proportionate to the risk class” and “appropriate for the type of device”. |  | reject. This is beyond the scope of the docuemnt.
annex I | CA6 | Annex I, PSUR | Ge | “ A summary of the results and conclusions of analysis of the PMS data collected ” It could be helpful to include description of when a PMS report should be written, how and when it should be updated and how it relates to the corresponding PMS plan, and its dates/timelines. |  | accepted
annex I | CA6 | Annex I, PMS plan | Ge | “ MDR/IVDR: Applicable for all device classes.” Note! This is not totally correct. E. g. for custom made devices there is no obligation for a technical documentation. |  | accept
5.2 | CA6 | 430 | Te | Comment: MDR and IVDR both state that PMS should be proactive. There is no requirement that PMS should also be reactive. However, excluding reactive data, e. g. complaints, would not make any sense. Thus, this guidance document should explain how also reactive data could be used in the PMS system. |  | accept
1 | NB1 | 91 | ge | Rephrase to add reference to MD legacy devices | Requirements for legacy devices have been detailed in MDCG 2021-21 for MDs and in MDCG 2022-8 for IVDs | accept only add wording for the legacy by referring to the MDCGs. Do
3 | NB1 | 148 | ge | The footnote n°4 does not refer to IVDR Which article should be referred to? |  | recital 80 MDR and recital 81 IVDR
3 | NB1 | 179 | ge | The reference on the IVDR annex is not Annex VIII but Annex XIII | Replace Annex VIII, Part B, Sections 4 and 5 of the IVDR by Annex XIII, Part B, Sections 4 and 5 of the IVDR | accept
PMS as an integral part o | NB1 | 213 | ge | The PMS data is not limited to what is listed in the Annex III 1 a) e.g. production data | […] the amount of data to be collected as inputs for the PMS processes including those listed in the Annex III 1 (a) | accept
4a | NB1 | 264 | ge | Add reference to IVDR | As referred to in Section 3 of Annex I MDR / IVDR | accept
4a | NB1 | 263 | ed | Typo, remove “the” | Risk analysis of the device should be covered in the PMS plan and be linked with […] | accept
5 | NB1 | 405 | ge | Wrong reference | Paragraph 5.1.2 | accept - will be reviewed when document is finalized
5 | NB1 | 405 | ge | Wrong reference | Paragraph 5.1.2 | accept
5.1.1 | NB1 | 414 | ge | On Information concerning serious incidents, including information from PSURs, and field safety corrective actions Trend reporting is non-applicable to serious incidents, article 88 MDR and 83 IVDR only cover non-serious incidents and undesirable side-effects Update of PSUR and SS(C P has to be considered for all data source, not only the data from serious incidents | Replace the whole cell How to utilize it by: Information concerning serious incidents or FSCA should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR. The purpose of this process is to detect any new risk or change in the frequency or severity of a known risk. When a change is detected, this information should be considered in the context of the overall risk assessment of the device, with due consideration given to any requirement to update the labelling, instructions for use, or other documentation associated with the medical device. | accept, text will be rewritten
5.1.1 | NB1 | 414 | ge | On Records referring to non-serious incidents and data on any undesirable side-effects The cell How to collect it does not provide any information on the collection process of such data | Replace the whole text by : There is an expectation that manufacturers have robust procedures to collect complaints and customer feedback, and evaluate if these complaints and customer feedback fulfil the criteria of an incident under Article 2(64) MDR/ Article 2(67) IVDR. Ref MDCG 2023-3 | accept
5.1.1 | NB1 | 414 | ge | On Records referring to non-serious incidents and data on any undesirable side-effects The cell How to utilize should provide more clarification about the detection of trends as defined in the Regulations | Replace the text by : The records referring to non-serious incidents and data on expected undesirable side-effects/ expected erroneous results should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. Non-serious incidents and expected undesirable side-effects/ expected erroneous results should be trended in comparison to pre-defined thresholds to identify and report any statistically significant increase in the frequency or severity of those incidents | partially accepted
5.1.1 | NB1 | 414 | ge | On Relevant specialist or technical literature, databases and/or registers The cell How to utilize : Information for PMS activities shall ensure that all thresholds and indicators are appropriate not only tresholds for trend reporting | Replace: This information should be used to update the risk management documentation and to ensure that thresholds (including those for trend reporting) are appropriate | reject , not the right place in the document.
5.1.1 | NB1 | 414 | ge | On Records referring to non-serious incidents and data on any undesirable side-effects The cell How to utilize should provide more clarification about the detection of trends as defined in the Regulations | Replace the text by : The records referring to non-serious incidents and data on expected undesirable side-effects/ expected erroneous results should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. Non-serious incidents and expected undesirable side-effects/ expected erroneous results should be trended in comparison to pre-defined thresholds to identify and report any statistically significant increase in the frequency or severity of those incidents | partially accepted. Same comment as in line 472
5.1.1 | NB1 | 414 | ge | On Relevant specialist or technical literature, databases and/or registers The cell How to utilize : Information for PMS activities shall ensure that all thresholds and indicators are appropriate not only tresholds for trend reporting | Replace: This information should be used to update the risk management documentation and to ensure that thresholds (including those for trend reporting) are appropriate | reject , not the right place in the document. Same comment as in line 474.
5.1.1 | NB1 | 419 | ge | On a proactive and systematic process… The cell Explanation: Information is not clear on how to organize the collection process while ISO/TR 20416 can provide some examples | Add: For each PMS data source, the manufacturer is expected to describe the approach taken to collect the data (see table 5.1.1). Example of data collection methods can be found in ISO/TR 20416. | rejected - ISO/TR 20416 is not a harmonized standard.
5.1.2 | NB1 | 419 | ge | On a proactive and systematic process… The cell Explanation: Information is not clear on how to organize the collection process while ISO/TR 20416 can provide some examples | Add: For each PMS data source, the manufacturer is expected to describe the approach taken to collect the data (see table 5.1.1). Example of data collection methods can be found in ISO/TR 20416. | tabel is revised
5.1.2 | NB1 | 419 | ge | On effective and appropriate methods and processes… The cell Explanation: Information is not clear on how to use the collected data. The goal of this continuous assessment of PMS data is also missing | Replace the whole text: For each PMS data source, the manufacturer is expected to describe the approach taken to analyse the data. Examples of data analysis methods can be found in ISO/TR 20416. In line with the MDR and IVDR, the methods employed should be proportionate to the risk of the device: For higher-risk devices, a more specific, focused PMS activity and a more detailed analysis of the resulting data may be required in order to generate conclusions. The PMS plan should specify the aims of these analysis methods and the type of information that is being sought (see table 5.1.1). | tabel is revised
5.1.2 | NB1 | 419 | ge | On suitable indicators and threshold values… The cell Explanation: the sentence is unclear, tresholds and indicators should be pre-defined and then re-evaluate with the conclusions from the continuous reassessment of the PMS data | Replace the whole text: Indicators and threshold values should be defined by the manufacturer based on available data (e.g. data from risk management, clinical evaluation/ performance evaluation, PMS, etc.). The data collected during PMS activities can be used for the re-evaluation of these indicators and threshold values. | tabel is revised
5.1.2 | NB1 | 419 | ge | On methods and protocols to manage the incidents… The cell Explanation: the 1st sentence is not applicable to this line but applicable for the previous one one “effective and appropriate methods and tools” | Move the sentence to the previous line: The post-market surveillance plan should describe the manufacturer’s processes and methods for investigating complaints and how to collect and analyse data related to market-related experience collected in the field. | tabel is revised
5.1.2 | NB1 | 419 | ge | On methods and protocols to manage the incidents… The cell Explanation: the documents does not bring any explanation on methodology | Add: The plan should describe the trending methodology, the observation period, the thresholds established to detect any statistically significant increase in the frequency or severity of non-serious incidents... | tabel is revised
5.1.2 | NB1 | 419 | ge | On PMPF plan as referred to in … The cell Explanation: the section explains that PMCF/PMPF plan is required in certain cases. This statement does not reflect what is described in the section 4 of the guidance and in the Regulations. PMCF/PMPF exemptions are applicable for limited devices | Replace with: PMCF/PMPF plan should cover both general and specific methods as specified in section 3 line 158-182 and section 4 line 321-324. The aims of the PMCF are covered in section 4 line 292 to 299. The guidance should provide examples where PMCF/PMPF is not applicable or not appropriate. | tabel is revised
5.1.1 | NB1 | 414 | ge | On Information from trend reporting The cell How to utilize : same remark as earlier the need of updating SSCP,labelling etc is applicable for all types of PMS data | Remove the sentence : The manufacturer should also consider the need to update the labelling and where applicable update to the SS(C)P to inform users of the device of any new or emerging risks. | accept. Sentence was removed.
5.1.1 | NB1 | 414 | ge | On Information from trend reporting The cell How to utilize : same remark as earlier the need of updating SSCP,labelling etc is applicable for all types of PMS data | Remove the sentence : The manufacturer should also consider the need to update the labelling and where applicable update to the SS(C)P to inform users of the device of any new or emerging risks. | accept. Sentence was removed. Same comment as in line 473.
5 | NB1 | 403 | ge | Wrong reference | In paragraph 5.1.1 | accept - will be reviewed when document is finalized
5 | NB1 | 403 | ge | Wrong reference | In paragraph 5.1.1 | accept
5.1.1 | NB2 | similar devices | GE | Based on our experience manufacturers do not understand that as similar devices the same ones available in TD 1. 1.2.b shall be used, to ensure the coherence of the TD and the coherence of the clinical/PMCF/PMS data. | Please add a clarification that the similar devices shall be the same as the ones declared in Part 1.2.b of the TD. | accept we will refer to the definition of mdcg 2020.6
5.1.2 | NB2 | p16 Table – lines 4-6 | TE | While the MDR is talking about methods, tools and/or protocols it is obvious that it shall be performed in the QMS of the manufacturer. | Please add clearly that the reference to the relevant procedure and/or instruction shall be added here. | accept
5.1.2 | NB2 | p16 Table Lines 1-3 | TE | We see as best practice that several manufacturers started to create tables for the first 3 requirements, namely data collection for their own device/for similar device / methods and processes to assess / indicators/threshold values We think it makes the activity and its description more transparent, we would suggest adding this option, explanation as well. | See the comment | reject. The purpose and aim of the guidance is to be a top level overarching guidance. This request is too detailed information for the aim of the document.
5.1.2 | NB2 | p16 2nd line of the table | TE | “Effective and appropriate methods and processes to assess the collected data;” is the title of this line – the explanation does not really cover the methods. The sentence “For example, a class III implant is likely to require specific PMCF activities and the methods of analyzing the associated data should be appropriate for its risk classification.” shall be deleted from here, it should be added to the PMCF plan requirements in the last line of the table. | Modify explanation, add the following: For each data required by Annex III. 1.b the detailed methods and processes required for collection and assessment of data shall be described. For example for literature search the methodology (data sources, search terms) and frequency of search including the person responsible for search, data approval and evaluation; for information from trend reporting: the frequency of evaluation, the considered time period etc. | partially accept.
annex I | NB2 | p18 Figure 2 | TE | In the Figure 2 Audits/Inspections are mentioned as inputs – please revise the figure, as audits are not mentioned as inputs neither in the MDR, nor in the guidance | Change to “Data of similar devices”, which would clarify the figure | Markus and Rachel to discuss was part of figure review.
annex I | NB2 | annex ! p18 5.3 | ED | Although the guidance contains that each PMS plan should include as the PSUR or PMS plan and its frequency as an output, but please make it more pronounced. | The post-market surveillance plan should also provide a direct reference to the output of plan, namely PSUR or PMS plan, including their frequency and to the procedures in relation to the output of the post-market surveillance data into a post-market surveillance report for class I devices (Article 85 MDR)/Class A &amp; B devices (Article 80 IVDR) and for a periodic safety update report for all other classifications per Article 86 MDR/Article 81 IVDR. | accept
annex I | NB2 | Annex Ip19 5.4 | ED | Conclusions need to be drawn up in a summary report. the PMS report or the PSUR8. / typo error | Conclusions need to be drawn up in a summary report (the PMS report or the PSUR) | accept
1 | NB3 | 122-123 | GE | repetition | Remove it | accept
2 | NB3 | 135-135 | Ed | Left over text? | Remove it | accept
3 | NB3 | 156 - 209 | GE | Structure of the document is not clear | Please reconsider the full structure of this part of the document. It is confusing to us now and seem to have a lot of repetition. | accept the need for revision
3 | NB3 | 171 - 182 | Te | This paragraph is causing confusion to the difference between the PMCF and the PMS. | Please explain that while PMS is the general and wide proactive collection of post-market information, PMCF is a pro-active process just focused on clinical / performance data – so in fact a part of the PMS activities including CER update. | accept the need for revision
3 | NB3 | 139 - 141 | TE | Discussion might come whether it is placed on the market or not. | Change placed on the market to when DoC is signed | reject
3 | NB3 | 149 | Ed | Footnote 4 is not clear. What is 80 in the MDR and where is the link to the IVDR? | No proposed change since it is not clear what is meant. | reference to th ecorrect articles
3 | NB3 | 150 | Ed | Remove the abbreviation PMS. Already mentioned | Remove (PMS) | accept
3 | NB3 | 156 | Ed | In simpler terms, indicates that the same is written down as what is written above in a more easy way. But this is not the same as what is written above. | Remove in simpler terms | accept
3 | NB3 | 160 - 165 | ED | Repetition of the paragraph below. | Please remove this paragraph and add 163 – 165 to section below 186. | accept
3 | NB3 | 193 | Ed | Remove the word and | Remove the word and | reject
PMS as an integral part o | NB3 | 213 | Te | Add more points | The manufacturer identifies a clear focus how they are accessing the PMS data. Which specific aspects require focus for the medical device? How will this information be filtered from all data collected? | make more practical
PMS as an integral part o | NB3 | 213 | Te | the amount of data to be collected as inputs | What is meant by this? Please specify in more detail. | resolved - see line 193
4a | NB3 | 262 | Te | Footnote 6, which guidance is that? |  | MDCG Guidance on Trend Reporting that is currently drafted. Footnote only to be included if Guidance oin Trend Reporting is published.
4a | NB3 | 264 | Ed | Add IVDR | (as referred to in Section 3 of Annex I MDR/IVDR). | accept
4c | NB3 | 311 | Ed | Missing and i | Identifying | accept
4c | NB3 | 330 - 331 | Ed | Wrong sentence | results in a PMCF/ PMPF evaluation report that shall be part of the clinical evaluation performance evaluation report. | Reject - sentence has been restrcutured.
4e | NB3 | 359 - 364 | Te | It is not sufficiently clear. | A more clear description that overall PMS data shall be analyzed for potential vigilance and required FSCAs. | accept
4f | NB3 | 372 | ed | Word missing | to identify if there are any | accept
4g | NB3 | 380 - 382 | Te | Please consider the other way around. The manufacturer shall consider if their device is impacted to the information which is found with regard to similar medical devices. | The manufacturer shall consider if their device is impacted to the information which is found with regard to similar medical devices. | reject-
4g | NB3 | 384 | Te | Add the impact on the risk management | appropriate actions would be taken to limit the impact to other users of the device and add update risk managements. | beyond the text - reject
5 | NB3 | 403 and 405 | Ed | 6.1.1 should be 5.1.1 | 6.1.1 should be 5.1.1 | accept
5.1.1 | NB3 | 414 (table). 3rd row | Te | Where would complaints be listed? In the 2nd or third row? Both sections mention complaints, but the how to use is different for both sections. | I would expect it to be in the second row. | accept
5.1.1 | NB3 | 414 (table). 8th row | Ed | Double, same as 2nd row | Remove this row | accept
5.1.2 | NB3 | 419 (2nd table) row 6 | Ed | 6th row, wrong footnote number and footnote is not clear. | Update text in footnote and change the number of the footnote to 7 | accept
5.1.2 | NB3 | 419 (2nd table) row 8 | Ed | Add IVDR reference | as described in Article 78, 79 and 81 IVDR and 83, 84 and 86 MDR, | accept
5.2 | NB3 | 427 | te | In that way | It is not clear in what way. | accept
5.2 | NB3 | 429 - 433 | Ed | repetition | Remove it here and leave it in in section 3 (around row 190). | accept
5.4 | NB3 | 481 | Ed | Wrong footnote number | Footnote number should be 8, not 9 (footnote itself) | accept
annex I | NB3 | 520, table 3, row 6 | Te | It states Clinical Evaluation Plan/ Performance Evaluation Plan. What is meant by this? | Please elaborate a bit more on this point to make clear what is expected here. | reject. Not in scope of this document.
annex I | NB3 | 520, table 3, row 9 | Ed | •a summary of the results and the conclusions of the analyses/assessment of the post-market surveillance data defined in the PMS plan. | •a summary of the results and the conclusions of the analyses/assessment of the post-market surveillance data resulting from the execution of the PMS plan. | Text was revised.
ge | NB4 |  |  | Use of must/shall etc throughout the PMS guidance: | As the present document is not deemed legally binding, the use of must / shall etc. should be avoided, if not occurring in extracts of the MDR/IVDR | accept
ge | NB4 |  |  | Punctuation need to be reviewed, as e.g. “.” are used where “,” seem to be indicated (e.g. 479, 482-485) |  | accept
2 | NB4 | 134-135 | Ed | Is this not an internal comment, that should be removed? |  | accept
3 | NB4 | 199 |  | Monitoring of claimed clinical benefits is not a requirement of MDR, Annex XIV, Part B | This includes the monitoring of known residual risks as well as risks not previous identified (emerging risks) for the patient or user. | accept the need for revision
3 | NB4 | 182 |  | “ A manufacturer does not need to wait…” | Wording “does not need to wait” to be reviewed | accept
3 | NB4 | 170 | Ed | Annex VIII is wrong, must be Annex XIII |  | accept
PMS as an integral part o | NB4 | 213 | Ed | End of line, parenthesis must be removed. |  | accept
4a | NB4 | 262 | Ed | Footnote 6 is not clear |  | accept
4a | NB4 | 264 | Ed | Missing within the parenthesis: “…and IVDR)” |  | accept
4c | NB4 | 292-299 |  | As the list of items are not providing guidance beyond what is written in the MDR, the wording should not be changed and should be in line with the MDR! e.g. “ensuring the continued acceptability of the benefit-risk ratio” | Update to be in line with regulation | accept
4c | NB4 | 285 | Ed | For IVDR, must be Art. 78.3.c instead of 78.3.b |  | accept
4c | NB4 | 301 | Ed | Clinically relevant must be changed to performance relevant |  | accept
4c | NB4 | 302 | Ed | Post-market clinical follow-up must be changed to post-market performance follow-up Clinical data must be changed by relevant scientific data |  | accept
4c | NB4 | 303 | Ed | Performance evaluation results must be changed by performance data |  | accept
4c | NB4 | 311 | Ed | Dentifying must be corrected to Identifying |  | accept
4c | NB4 | 330 | Ed | End of the line, r/ must be corrected to / |  | Reject - sentence has been restrcutured.
4d | NB4 | 350 | Ed | Clinical evaluation report must be corrected to clinical/performance evaluation report |  | accept
4d | NB4 | 351 | Ed | A period (dot) is missing after “report” |  | accept
4e | NB4 | 362 |  | Crosscheck with MDR art. 2 (68); description of FSCA to be in line with definition in Art. 2 | ‘field safety corrective action’ means corrective action taken by a manufacturer for technical or medical reasons to prevent or reduce the risk of a serious incident in relation to a device made available on the market; | accept
4f | NB4 | 372 |  | “ The information should be considered to identify if there any options to improve…” To add “are” | The information should be considered to identify if there are any options to improve | accept
4f | NB4 | 379 |  | “… share the same intended purpose or potentially similar design characteristics” | Or ? should it not read and? | partially accepted - modify to and/or
4g | NB4 | 382-384 |  |  | “where this could happen, appropriate actions would be taken to limit …” As this is case dependent, appropriate actions could not be indicted. To be clarified what is expected here and in which document? | accept paragraphs are rephrased
5 | NB4 | 400 | Ed | Not clear what is meant by Figure 1; not such legend could be found in the document. |  | accept
5 | NB4 | 403 | Ed | 6.1.1 must be changed to 5.1.1 |  | accept
5 | NB4 | 405 | Ed | 6.1.2 must be changed to 5.1.2 |  | accept
5.1 | NB4 | 406 |  | 5.1 Planning: To be clarified whether a PMS plan can be of a general nature or need to be device related | To be clarified whether a PMS plan can be of a general nature or need to be device related | clarified
5.1.1 | NB4 | 414 |  | e.g. 5.1.1 “ Relevant specialist or technical literature, databases and/or registers” p. 12: related to screening | add for screening that search terms are to be specified | reject - too detailed for a PMS Guidance
5.1.1 | NB4 | 414 |  | Table: Any reason for changing the order compared to Annex III, 1(a) | To keep the order as in Annex III, 1(a) | accept
5.1.1 | NB4 | 414 |  | First line “…plans a survey…” | Would this be considered part of a PMCF? | accept paragraphs are rephrased
5.1.1 | NB4 | 414, p. 12 first line |  | “… should encourage reporters…” | Not clear, which kind of encouragement could be initiated and which would be deemed acceptable / allowed? | accepted - refer to Vigilance Guidance
5.1.1 | NB4 | 414, p. 12 second line |  | Right column “… to notify Competent Authorities in relation to articles 88 …” | Art. 88 does not include reporting to CAs but using Eudamed left column refers to Art. 88 only, right column refers to MDR Art. 89 in addition. | accept
5.1.1 | NB4 | 418, p. 14 first line |  | Right column “… conducting a benchmarking exercise against…” | Bench marking exercise? | accept, text will be mofified
5.1.1 | NB4 | 414 (Table) | Ed | The second line on the table (Information including feedback…) is repeated twice; it is found below again although the text in 2nd and 3rd columns is slightly different. |  | accept
5.1.1 | NB4 | 414 (Table) | Ed | In the 6th line (Relevant specialist or …), 3rd column, and class C and class D devices. must be added at the end. |  | accept
5.1.2 | NB4 | 418, p. 14 second line |  | “… the risk of the device…” Should read risk class? | “… the risk class of the device…” | accept
5.1.2 | NB4 | 418, p. 14 5th line |  | MDCG Q&A trend | MDCG 2023-3 | reject - MDCG 2023-3 is a Q&A on Vigilance, not Trend Reporting.
5.1.2 | NB4 | 418, p. 15 first line |  | IVDR articles, MDR missing |  | accept
5.1.2 | NB4 | 418, p. 15 first line |  | check: Is in “including the PMS plan” correct, as references to the procedures are part of the PMS plan , see 5.3 below |  | accept, reworded.
5.1.2 | NB4 | 418, p.14 Last line |  | “…, for submitting trend reports to competent authorities…” ? Crosscheck with MDR requirements |  | accept
5.1.2 | NB4 | 419 (table) | Ed | In the 6th line, second column, last part of the text is repeated and the reference number does not match with the one found at the page foot. |  | accept
5.1.2 | NB4 | 419 (table) | Ed | The reference for Footnote 9 is not found in the table text; presumably it should correspond to line 8, second column. |  | accept
figure 437 | NB4 | 437 (diagram) | Ge | The diagram is in fact overlapping with information in table 5.1.1; I would include in the foot note: Figure 2: Examples of PMS data and their respective action types (not exhaustive; see table 5.1.1 above). |  | accept, figure is deleted
5.3 | NB4 | 462 | Ed | Not clear what reference to “Annex IIIb” means. Perhaps Annex III, 1.b. 8th indent. |  | accept
5.4 | NB4 | 470 |  | “…. maintain a post-market surveillance system for each device as an…” | PMS System for each device or applied for each device? | accept, figure is deleted
5.4 | NB4 | 476-477 |  | “… for collecting information. The data needs to..” One sentence? | for collecting information, the data needs to | accept
5.4 | NB4 | 467 | Ed | Chapter 6 must be changed by Chapter 4 |  | accept
5.4 | NB4 | 476 | Ed | So besides planning the PMS, and follow the procedures for collecting information. The data needs to be analysed and assessed. To: | So, besides planning the PMS and follow the procedures for collecting information, the data needs to be analysed and assessed. | accept
5.4 | NB4 | 479 | Ed | End of line is incorrect. Proposal: …in a summary report (PMS report or the PSUR). The reference number (8) does not match with the footnote number (9). |  | accept
5.4 | NB4 | 482 | Ed | Sentence must start with capital letter. The entire paragraph needs capital letters after the periods (dots); revise in general use of periods in this paragraph. |  | accept
6 | NB4 | 498 |  | “some kind of documentation”? Wording to be reviewed |  | accept
6 | NB4 | 494 |  | In that context, it shall report it to be replace by the manufacture | In that context, the manufacturer shall report | accept
annex I | NB4 | Annex I . p. 21 |  | NB evaluate and add its evaluation report in EUDAMED. A report is not required acc. to MDR Art 86.2 Add its evaluation to Eudamed | NB evaluate and add its evaluation to EUDAMED. | accepted - removed
annex I | NB4 | Annex I, p. 18 |  | Allows update of the technical documentation. Wording “allows” to be reviewed |  | accepted. Wording has been revised.
annex I | NB4 | 529 (table) | Ed | First line (header), Articles of the regulations can be removed. |  | accept
intro | NB5 | 71-73 | GE | The use of the term “mainly” is misleading. The IVDR says that compliant devices are under the sole responsibility of the manufacturer; refer to Annex IV (2): 2. A statement that the EU declaration of conformity is issued under the sole responsibility of the manufacturer; Further, for Class A devices IVDR or Class I MDR no NB5 involvement is required, refer to Praeamble IVDR (56) The conformity assessment procedure for class A devices should be carried out, as a general rule, under the sole responsibility of manufacturers, since such devices pose a low risk to patients. For class B, class C and class D devices, an appropriate level of involvement of a notified body should be compulsory | Remove the word “mainly”: It is essential to understand that within the European legal framework the responsibility for placing a safe medical device rest with the manufacturer of the device. | Accept - that part of the sentence was removed as it pertains to 'placing on the market' and not PMS
intro | NB5 | 74-77 | TE | Suggest to take into account the definition of PMS in Article 2 (63) ‘post-market surveillance’ means all activities carried out by manufacturers in cooperation with other economic operators to institute and keep up to date a systematic procedure to proactively collect and review experience gained from devices they place on the market, make available on the market or put into service for the purpose of identifying any need to immediately apply any necessary corrective or preventive actions. | Rephrase the sentence as follows: Once placed on the market, European regulations oblige the manufacturer to perform post-market surveillance (PMS) for the medical device to identify any need for any necessary corrective and preventive actions and to monitor whether the medical device continues to meet the general safety and performance requirements specified in the regulation. | see above
3 | NB5 | 170-172 | TE | Annex XIV, Part B, Sections 5 and 6 MDR/Annex VIII, Part B, Sections 4 and 5 of the IVDR emphasize the role of a manufacturer to actively monitor information obtained from post-market experience. | Correct the reference to IVDR: Annex XIII, Part B instead of Annex VIII, Part B | accept
3 | NB5 | 183 | TE | A manufacturer does not need to wait for information to be sent to them (e.g. via customer complaints) but can “proactively” collect and (…) “but can” is a bit misleading, as the entire concept is to have proactive activities, and not a voluntary “can” to do so | Replace “but can” with “but is expected: A manufacturer does not need to wait for information to be sent to them (e.g. via customer complaints) but is expected to “proactively” collect and (…) | accept
PMS as an integral part o | NB5 | 212-223 | TE | Annex III uses the word “shall” and not “must”. | Suggestion to align to Annex III and replace all “must” with “shall” | accept
4c | NB5 | 284-285 | TE | c) Update the clinical evaluation (Art. 61 MDR) / performance evaluation (Art. 56 IVDR) - Art. 83. 3 (b) MDR / Art. 78.3 (b) IVDR | Correct the reference to MDR /IVDR: Art. 83. 3 (c) MDR / Art. 78.3 (c) IVDR instead of Art. 83. 3 (b) MDR / Art. 78.3 (b) IVDR | accept
5.1.2 | NB5 | 5.1.2 | TE | Line effective tools to trace and identify devices for which corrective actions might be necessary; and A description of the process to identify product and product status throughout product realization should be provided | Suggest to replace “product” with “device” | accept
5.1.2 | NB5 | Table 5.1.1 | TE | Section: Information, including feedback and complaints, provided by users, distributors and importers | This section is repetitive and needs to be removed once in table ideally the first one according to MDR/IVDR annex III requirements | accept
5.1.2 | NB5 | Table 5.1.1 | TE | Section Information concerning serious incidents, including information from PSURs, and field safety corrective actions The following statement is unclear: Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. Trend reporting as per MDR art.88 and IVDR art. 83 refers to non-serious incidents, so it is not clear why it is referenced here. | Clarify the reference to trend reporting in the serous incident section | accept
5.3 | NB5 | 457 | TE | These procedures should consider the frequency and trigger points that may require an update to these reports and potentially further updates to the post-market surveillance plan (Article 85 / Article 79). | Based on the sentence, reference article 80 of the IVDR | accept should be 79
5.4 | NB5 | 465-468 | TE | Articles 83 (3) MDR and 78(3) IVDR state that the data collected by the manufacturer's post-market surveillance system shall in particular be used for input and to update processes in the system, as explained in chapter 6. The sentence above is referring to processes described in chapter 6 but chapter 6 is covering custom-made devices only. | Suggestion to rephrase “Articles 83 (3) MDR and 78(3) IVDR state that the data collected by the manufacturer's post-market surveillance system shall in particular be used for input and to update processes in the system to ensure Annex I compliance. | accEpt
5.4 | NB5 | 477-486 | ED | Reformat the following paragraphs for clarity: So besides planning the PMS, and follow the procedures for collecting information. The data needs to be analysed and assessed. The findings shall be used to assess whether it is necessary to adjust the device, to improve the usability, performance and safety of the device. Conclusions need to be drawn up in a summary report. the PMS report or the PSUR8. to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes. | · Suggested format: So besides planning the PMS and apply the procedures to collect information, the manufacturer needs to analyse and assess the data. The findings shall used to assess whether it is necessary to adjust the device, improve the usability, the performance and safety of the device. Conclusions need to be drawn up in a summary report, either the PMS report or the PSUR. To start the next period of PMS, a review of the PMS plan is necessary. The scope of this review is to determine if: the plan is still up to date the data sources used are still appropriate the information collected address the objectives of the PMS plan the findings are used for the relevant processes. | accept
ge | NB6 |  | te | A common question for IVDs not addressed in the guidance: Does the PMS plan have to be specific for a device or can it be for a group of devices such as all Class B or all cancer devices?
ge | NB6 |  | te | Clarification for Art.80 on appropriate timeframe for preparing a PMS report. |  | accept
3 | NB6 | Line 180 |  | Include information published in clinical guidelines regarding the use of the device or presented at clinical/professional society meetings | “…both through general PMCF/PMPF methods such as gathering of clinical experience (e.g. via general surveys or clinical conference presentations) or screening of scientific literature/clinical guidelines or through specific methods such as evaluation of suitable registers or PMCF/PMPF studies. | accept the need for revision
3 | NB6 | Line 202/203 |  | include reference to actions taken under 'Preventative Health Protection measures' in line with Article 98 (MDR) | “Relevant data and information gathered through post-market surveillance, as well as lessons learned from any implemented preventive and/or corrective actions, or actions taken under Article 98, should be used to update any relevant part of technical documentation, such as those relating to risk assessment and clinical evaluation.” | accept the need for revision
4a | NB6 | Line 248 | Ed | typo -repetition of 'the' |  | accept
4a | NB6 | Line 263 | Ed | typo - excess 'the' in sentence |  | accept
4c | NB6 | 314 | te | No reference to identifying possible systematic misuse | Add ‘ identifying possible systematic misuse’ | Reject, it is already listed in the text.
4c | NB6 | Line 323 |  | Addition to include “published clinical advice about use of the type of device” e.g. medical device alerts published by CAs | “These general and specific methods and procedures of PMCF/PMPF should include post-market activities, such as gathering of clinical experience gained, feedback from users, screening of scientific literature, published clinical advice about use of the type of device, evaluation of suitable registers or PMCF/PMPF studies. | Reject. TF is of the opinion to stay away from repeating what has already been described in other MDCG guidance. Therefore, document will refer to already establied MDCG guidance. This guidance will not explain in detail how to update the clinical evaluation.
4c | NB6 | 302 | ed | For IVDR refers to ‘post-market clinical follow-up data’ | Change to ‘post-market performance follow-up data’ | accept
4d | NB6 | 350 | ed | States ‘clinical evaluation report’ | Change to ‘ clinical/performance evaluation report’ | accept
4e | NB6 | 362 to 364 | te | Refers to only one type of serious incident ‘ to reduce a risk of death or serious deterioration in the state of health associated with the use of medical devices’ | Change to ‘ to reduce the risk of recurrence of a serious incident associated with the use of medical devices’ | accept -sentence changed in alternative way
5.1.1 | NB6 | Page 12 | ed | Column 3 for relevant specialist or technical literature states ‘ The Clinical Evaluation Report’ | Change to ‘ The Clinical/Performance Evaluation Report | accept
5.1.2 | NB6 | Page 15 | te | Column 3 for PMPF plan refers to equivalence | Provide a description of what equivalence means for IVDs | reject - Not for PMS guidance. Only 2022-2 refers to equivalence of IVD.
5.1.2 | NB6 | Section 5.1.1 |  | Table row "Information concerning serious incidents, including information from PSURs, and field safety corrective Actions" could include reference to actions taken by regulators | “ Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. This information along with any actions/alerts issued by regulators should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR.” | reject
figure 437 | NB6 | Line 437 |  | - Include: ‘clinical guidelines’ in the literature review bubble Actions/advice issued by regulators regarding the use of the device in the blue bubble |  | accept, figure is deleted
ge | CA7 | 0 | Ed | Check punction throughout the document, as in some places . is missing or are double. | - | accept
intro | CA7 | 72 | ge | Sentence incomplete | Add “on the market” at the end of the sentence | accept - wording removed
1 | CA7 | 102 | Ge | Add that PMS should be part of the QMS system, to better connect to the next sentence | Add “Including a PMS system (sub i)” | accept
1 | CA7 | 115-116 and 121-122 | Ed | It is mentioned several times that PMS is part of the QMS system .This is also part of the paragraph 98-110. | Delete sentence 115-116 and 121-122 | see comment on line 73 above
1 | CA7 | 111 | Ed | Improve language | Replace “to be found” with “are laid down” | accept
2 | CA7 | 134 | Ge | Is using “we” appropriate in a MDCG guidance document? | Rephrase sentence to delete “we” | resolved - sentence deleted
3 | CA7 | 177 | Te | Emphasize that the manufacturer should be actively checking for available data, t | Add at the end of the sentence; “and to actively check for available information”. | accept
3 | CA7 | 182 | Ge | Rephrase to emphasize the proactive approach of PMS | Replace “does not need to” by “should not” | accept
3 | CA7 | 183 | Ge | Rephrase to emphasize the proactive approach of PMS | Replace “can” by “should” | accept
3 | CA7 | 190 | Ge | Cooperation with the CA’s might not be necessary | Add “, if needed” at the end of the sentence | accept
3 | CA7 | 166 | Ed | Delete “post market” from the end of the sentence, as this is repeated in the next part of the sentences | Delete “post market” from the end of the sentence | accept
PMS as an integral part o | CA7 | 212 | Te | It is not only about identifying potential sources, but more about using the appropriate sources. It should be prevented that the manufacturer has to use all potential sources, which are many as can be seen from CEN ISO/TR 20416 | Rephrase sentence: must identify potential data sources and select the appropriate data sources that the required type of data and amount of data is obtained, considering the risk class and type of device (see also Annex A of ISO CEN/TR 20416). From the selected data sources, data must be collected as inputs for the PMS processes……. | not fully understood; no need to refer to TR
PMS as an integral part o | CA7 | 207-208 | Ed | This text is superfluous and does not relate to the next paragraph | Delete line 207-208 | accept
4c | CA7 | 330 | Te | The PMCF report does not have to be part of the CER/PER, but shall be included in those reports, which leaves the manufacturer some options how to include those. | Replace “part of” by “included in” | Partly accepted, wording has been changed to "used to update"
4c | CA7 | 311 | ed | Missing letter: I | Change dentifying to Identifying | accept
4d | CA7 | 343 | Ed | Add performance at the end of the sentence for IVDs | Add “/performance” at the end of line 343. | reject already in th esentence
4f | CA7 | 373 | Ed | Improve language | to identify if there are any options | accept
4h | CA7 | 391 | Ge | Superfluous wording | Delete “ within the quality management system of the manufacturer” | accept
5 | CA7 | 403/405 | Ed | Correct references | Replace 6.1.1 and 6.1.2 with 5.1.1 and 5.1.2. | accept
5.1.1 | CA7 | 415, last row of table, 2nd column | Ge | “May wish to use” is open ended. Make more strict. | Replace “may with to” with “can” or “should” | accept
5.1.1 | CA7 | 414 (2nd row, 2nd column) | Ed | Use language used previously in paragraph | Replace “ market-related experience collected in the field” with “feedback” | reject
5.1.2 | CA7 | 419, 4th row, 2nd column | Ge | Do we really want to set thresholds based on the PMS results. That could lead to setting the thresholds so that the data are acceptable. Thresholds should be established before the analysis of the data. | Replace “The data obtained from these activities should also be used to set thresholds and indicators that will allow” with “Thresholds should be established before analysis of the data, allowing”. | accept, text modified.
7 | CA7 | 525 | Ed | World Health Organization was not involved in developing CEN ISO/TR 20416. | Delete “World Health Organization” | reject
4c | CA7 | 284 | ed | The references to the articles 83.3 (b) MDR and 78.3 (b) are incorrect. | Change the text to: Update the clinical evaluation (Art. 61 MDR) / performance evaluation (Art. 56 IVDR) - Art. 83. 3 (c) MDR / Art. 78.3 (c) IVDR | accept
4e | CA7 | 358-359 | Ed | Improve language | Replace “to determining, implementing and monitoring” with “ to determine, implement and monitor” | not relevant -text is changed
ge | CA8 |  |  | Use the word ‘shall’ instead of ‘should’ to be aligned with the regulation’s wording | ‘shall’ instead of ‘should’ throughout the whole document | accept, we will use 'must'
1 | CA8 | 91 |  | Legacy devices (already explained in another MDCG document) are using the article 120 to be placed on the market after May 2021. Legacy devices are certified under MDD/IVD but not according to the article 120(3)… | Delete the mentioned wording “…according to Article 120(3) MDR and Article 110(3) IVDR” | accept propose to only add wording for the legacy by referring to the MDCGs - Hanim checking for exisiting wording on legacy devices
3 | CA8 | 165-168 |  | The wording could be misunderstood as ‘the PMS is intended for monitoring the pre-market fase…’ | The PMS system is intended to continuously monitor the results of the pre-market phase (also known as development phase) with “real world” information from the post-market data and experience gained during the post-market phase and to identify and initiate appropriate measures or update the design and manufacturing when needed. | accept
4h | CA8 | 389-390 |  | Add condition for ‘when to not use trend report’. There are a few conditions for reporting “expected undesirable side-effects (MDR) or expected erroneous results (IVDR)” using “the trend report”. The conditions are: clearly documented … and quantified in the product/technical information … and are subjected to trend reporting. The 3. condition: The expected side-effects/ erroneous results must not exceed the threshold for determination of a serious incident, because Trend reporting requirements are for “incident” only. (Any side-effects could/may first be subjected/considered for the conditions of a ‘serious incident’ determination). | expected undesirable side-effects (MDR) or expected erroneous results (IVDR) which are not classified as “serious incident” | Reject -discussed in trend report Q&A
5.1.1 | CA8 | 414 |  | Reporting requirements of serious incidents are stated in article 87 of the MDR and article 82 of the IVDR. | “…Information relating to serious incidents should be used for trend reporting pursuant to article 88 87 of the MDR and article 83 82 of the IVDR…” | accept
5.2 | CA8 | 429 |  | Change the wording ‘express an expectation’ to ‘require’ | The regulations require that the post-market surveillance plan … | accept
annex I | CA8 | 529 |  | Add ‘serious incident’. | Describes the methods and protocols to monitor trends, to identify significant increase in frequency or severity of incidents and serious incidents. | reject. The text in the guidance is in line with the Regulations.
intro | CA9 | After 77 | ED | Add list of explanations of all abbreviations used in the text of the manual | MDR IVDR MDCG PMCF PMPF PMS PSUR QMS SS(C)P FSCA EEA | reject
1 | CA9 | 92 | GE | We propose to add “guidance document” before MDCG 2022-8 | Replace “MDCG 2022-8” with “guidance document MDCG 2022-8” | accept propose to only add wording for the legacy by referring to the MDCGs
1 | CA9 | 104 | ED | Add “to” to wording of “appropriate the type of device” | Appropriate to the type of device | reject (see above)
1 | CA9 | 108 | ED | Abbreviation “PMCF” shall be extended at first appearance in the text. | Add” Post-market Clinical Follow-up” before PMCF abbreviation | accept
1 | CA9 | 109 | ED | Add abbreviation “PMS” after first appearance in the text of “Post-market surveillance” |  | accept - added in introduction
1 | CA9 | 111 | ED | Add “are” to “IVDR requirements for the risk management are to be found” |  | resolved see above
1 | CA9 | 112 | ED | Lower case in wording “Performance evaluation” | Replace with “performance evaluation” | accept
3 | CA9 | 158 | ED | Placed on the market instead of “plced in the market” (as in text) |  | accept
3 | CA9 | 162 - 164 | ED | There is a need to align plural or singular grammar form in sentence: “It is vital to understand that a manufacturer needs to demonstrate how he they observes and checks the safety and performance of the medical device and keeps it under systematic review once it has been placed on the market.” or in plural: “ It is vital to understand that manufacturers needs to demonstrate how they observe and check the safety and performance of the medical device and keep it under systematic review once it has been placed on the market.” |  | accept
3 | CA9 | 172 | ED | There is a need to give full name of abbreviation PMPF – “Post-market Performance Follow-up” as it appears here for the first time |  | accept
PMS as an integral part o | CA9 | 224 | GE | Add “side”- to “effects” in wording of: “ corrective actions if side-effects are identified that could have a significant impact on the benefit-risk analysis” |  | accept text revised
PMS as an integral part o | CA9 | 213 | ED | Remove open bracket “(“ from the end of the line |  | accept
4 | CA9 | 237 | ED | We propose plural form of verb “to request” as it refers to both articles mentioned earlier like ”Article 83 MDR and Article 78 IVDR specifically requests that information generated from..” |  | accept
grafic 1 | CA9 | 241 | ED | “SS(C)P” – abbreviation needs to be explained |  | agreed, graphic will be taken out or replaced with a different graphical representation.
4a | CA9 | 243, 271, 284, 340 | ED | Add “of” to the title of chapter like: “Update of the benefit-risk determination and to improve the risk management – Art. 83.3 (a) MDR / Art. 78.3(a) IVDR” Or Rather use continuous form “updating of” in the titles a,b,c, d to stress the fact of continuity of such manufacturer`s obligation Writing like: “Updating of the benefit-risk determination and to improve the risk management…” |  | accept
4a | CA9 | 263 | ED | Cancel one “the” in: “and be the linked with the” |  | accept
4b | CA9 | 275-276 | ED | Data are in plural form in MDR “post-market surveillance data is are considered” |  | accept
4b | CA9 | 279 | ED | Maybe align contradiction with previous word side-effects to plural form like: “ side-effects and contraindications” |  | accept
4c | CA9 | 284 | ED | See comment on line 243 (line 264 in the Exccel spreadhseet) |  | accept
4c | CA9 | 308 | ED | Note no 6 misspelling “trendreprot” instead of “trend report” |  | accept
4c | CA9 | 311 | ED | “dentifying and analysing emergent risks” replace with “identyfying and analysing emerging risks” |  | Reject -emergent is stated in the regulation
4c | CA9 | 330 | ED | Cancel “r/” at the end of the line |  | Reject - sentence has been restrcutured.
4d | CA9 | 350.351 | GE | We presume that reference should be to “Periodic Safety Update Report (PSUR)” instead of “Periodic Summary Update report”? |  | accept
4d | CA9 | 340 | ED | See comment on line 243 (line 264 in the Exccel spreadhseet) |  | accept
4e | CA9 | 356.357 |  | We have to keep one wording scheme, if we wrote “analyse” we should replace “utilize” with “utilise” |  | accept
4f | CA9 | 366 | ED | See comment to line 354 |  | accept
4g | CA9 | 384 | ED | utilise instead of “utilize” |  | accept
5 | CA9 | 404 | ED | utilised (as in MDR text) instead of “utilized” |  | accept
5.1 | CA9 | 408 | ED | Cancel “for” in: “ based on the requirements for whicha are set out in Section 1 of Annex III” |  | accept
5.1 | CA9 | 410 | ED | utilised instead of “utilized” |  | accept
5.1 | CA9 | 411 | ED | Align verb with noun in singular or plural form (“manufacturer has” or “manufacturers have”) “This supports the manufacturer in demonstrating that they have he has in place an…” |  | accept
5.1.1 | CA9 | 414 3rd row 3rd column | GE | When trends are identified or there is evidence of a new or emerging risk, consideration should be given to whether field safety corrective actions are required, which could include the publication of a field safety notice. We propose to change “actions” from plural to singular form “action” as later we write about FSN in singular |  | accept
5.1.1 | CA9 | 414 3rd row 3rd column | GE | Text: “Information relating to serious incidents should be used for trend reporting pursuant to article 88 of the MDR and article 83 of the IVDR. This information should also be analysed in relation to article 89 of the MDR and article 84 of the IVDR.” Is in contradction to article 88 which states clearly that: „Manufacturers shall report, by means of the electronic system referred to in Article 92, any statistically significant increase in the frequency or severity of incidents that are not serious incidents” In our opinion trend reporting referred to in art. 88 cannot be applied to serious incidents. We propose to remove this indent as a whole. |  | accept
5.1.1 | CA9 | 413 | ED | Maybe rather “utilise” in place of “utilize” as MDR rather opts for UK spelling vs. US (however in onea place MDR uses “utilize”) |  | accept
5.1.1 | CA9 | 414 first row of table 3rd column | ED | Replace How to utilize it with “How to utilise it” rationale: as in the comment to line 413 |  | accept
5.1.1 | CA9 | 414 8th row 2nd column | ED | Change to plural form of “authority” as there are plural websites referred to later in the sentence “Competent Authorityies’ websites and databases, including those outside of the EEA” |  | accept
5.1.1 | CA9 | 414 8th row 3rd column | ED | Change “is” to “are” as data are used in plural form in MDR text “The data that is are gathered via the activities listed within Annex III 1.1(a)” |  | reject
5.1.2 | CA9 | 419 Row 6 Column 2 | GE | We propose to replace the proposed text with “description of methods and protocols aimed at analysis of all incidents to identify serious incidents (reportable) and incidents (non reportable but being subject to trend analysis and reporting accordingly to art. 88 MDR” |  | reject - however, we have modified the text.
5.1.2 | CA9 | 419, row 5 column 1 and 2 | GE | We do not understand explanation of MDR text: — effective and appropriate methods and tools to investigate complaints and analyse market-related experience collected in the field; Given as: “Market related experience, e.g; experience with the medical device technology; experience of the medical device experience of the medical device within a specific situation (e.g., hospital use, home use, or new patient group).or ergonomic features” Whose “experience: of device”- users or MF? What experience has in common with tools and methods of investigation? We propose to move here from column 2 row 6 description: The post-market surveillance plan should describe the manufacturer’s processes and methods for investigating complaints and how to collect and analyse data related to market-related experience collected in the field |  | accept
5.1.2 | CA9 | 414 5th row 3rd column | GE | Add: „The manufacturer shall specify how to manage the non-serious incidents and the methodology used for determining any statistically significant increase in the frequency or severity of such incidents, as well as the observation period, in the post-market surveillance plan” (citation from MDR art. 88) In line preceding: “The outputs of trend reporting should be used to notify Competent Authorities in relation to articles 88 and 89 of the MDR and articles 83 and 84 of the IVDR. The manufacturer should consider whether field safety corrective actions are required to be issued including field safety notices. The manufacturer should also consider the need to update the labelling and where applicable update to the SS(C)P to inform users of the device of any new or emerging risks.” |  | noted - Trending TF will be consulted
5.1.2 | CA9 | 414 5th row second column | GE | Replace: “ Articles 88 MDR and article 83 IVDR require manufacturers to trend (?) and report any statistically significant increase in the frequency or severity of non-serious incidents, or undesirable side effects.” With: “ Articles 88 MDR and article 83 IVDR require manufacturers to report any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side-effects that could have a significant impact on the benefit-risk analysis” (as in MDR) |  | accept
5.1.2 | CA9 | 414 6th row 3rd column | GE | Add „update” at the end of: „This information should be used to help determine whether there are any new or emerging risks associated with the device. The outputs of screening the literature and any registries should be considered as inputs to the clinical evaluation process and the overall benefit-risk assessment update” |  | reject
5.1.2 | CA9 | 419 Row 11 Colum 2 | ED | Change form of actions descritpion from infinitve to noun “ There are many reasons why a PMCF/PMPF plan may be required these could include: Identify residual risk Demonstrate that novel features do not impose any increased risk Collect clinical or performance data to support the lifetime of the device Confirm the safety and performance profile of the device where the initial approval was based on equivalence” to: “ There are many reasons why a PMCF/PMPF plan may be required these could include: Identification of residual risk Demonstration that novel features do not impose any increased risk Collection of clinical or performance data to support the lifetime of the device proper estimation Confirmation of the safety and performance profile of the device where the initial approval was based on equivalence” |  | accept
5.1.2 | CA9 | 419 row 3 column 2 | ED | Remove surplus dot at the end “for its risk classification. .” |  | accept
5.1.2 | CA9 | 419 Row 8 Column 2 | ED | Replace “its analysis” with “their analysis” as data referred to in this sentence are plural noun in MDR. the data to be collected and recorded, its their analysis and the requirements for the PMS report / PSUR |  | accept
5.4 | CA9 | 482 | ge | We propose to change: “to start the next period of PMS, a review of the PMS plan is necessary. this review is to determine if the plan is still up to date, if the proper data sources are still in use. if the collected information address the objectives of the PMS plan. and whether the findings are used for the relevant processes.” With: For manufacturers of class I devices after preparing of a PMS report or for manufacturers of class IIa, IIb, III devices after preparing of PSUR, to start the next period of PMS, a review of the PMS plan is necessary. This review is to determine if the plan is still up to date, if the proper data sources are still in use, if the collected information address the objectives of the PMS plan, and whether the findings are used for the relevant processes” |  | accept
5.4 | CA9 | 480 | ed | Change: “PSUR8” with “PSUR” |  | accept
5.4 | CA9 | 476-477 | ED | We propose to change: “So besides planning the PMS, and follow the procedures for collecting information. The data needs to be analysed and assessed.” To: “Therefore besides planning the PMS, and following the procedures for collecting information, the data need to be analysed and assessed.” |  | accept
annex I | CA9 | 529 Row 3 Column 2 | ed | Replace “Analyzing relevant data” With “Analysing relevant data” (as in art. 83.2 MDR) |  | accept
1 | CA9 | 98, 99 | GE | We propose that after first full citation like Medical Devices Regulation 2017/745 (MDR) to stop using the extended names of regulations beside commonly used and accepted abbreviations | Replace “Medical Device Regulation (EU) 2017/745 (MDR)” with simply “MDR” and similarly “ In Vitro Diagnostics Regulation (EU) 2017/746 (IVDR)” with “IVDR” only | accept
3 | CA9 | 160 | ED | Add “that” in wording of “to ensure that the medical device continues” |  | accept
4a | CA9 | 258 | ED | Word “data” in MDR are always plural so we propose to amend wording as follows “the collected data has have any impact on…” |  | accept
4b | CA9 | 278 | ED | Add “of” in the “and monitoring of the identified side-effects” |  | accept
4d | CA9 | 347 | ED | Add point after “g” in e.g. |  | accept
4e | CA9 | 354 | ED | We propose, like in previous titles (a-d) see comment to line 243, to use continuous form like: “Identifying of the needs for preventative, corrective, or field safety corrective actions” |  | accept
4g | CA9 | 376 | ED | See comment to line 354 |  | accept
4h | CA9 | 387 | ED | See comment to line 354 |  | accept
5.1.1 | CA9 | 414 second row second column | ED | 1. the methods and tools that that are used to investigate complaints - remove one “that” (duplicated) 2. Manufacturers are expected to - collect feedback Remove “-“ |  | accept
5.1.1 | CA9 | 414 7th row 3rd column | ED | Chnge „is” to „are” in reference to data as data are used in MDR text as plural Some post-market surveillance data, including certain types of complaints, may be considered to be clinical data (MDR) or performance data (IVDR) and is are therefore a relevant input to the clinical evaluation of the device. | ac | accept
5.1.2 | CA9 | 419 Row 10 Column2 | ED | Replace “realization” with “realisation” (MDR uses UK spelling of this word) |  | accept
3 |  | 185, Footer page 5 | te | Footer page 5 (80) MDR seem to be missing IVDR information | If applicable, reference to applicable IVDR articles should be added | accept