# Interplay between the Medical Devices Regulation (MDR) & In vitro Diagnostic Medical Devices Regulation (IVDR) and the Artificial Intelligence Act (AIA)

Introduction

This document provides a first set of answers to the most frequently asked questions related to the joint application of the AIA and the MDR or IVDR. The primary aim is to facilitate the application of those regulations by medical device manufacturers, notified bodies and competent authorities. All references to ‘manufacturer’ within the meaning of the MDR/IVDR should be understood as references to ‘provider’ in accordance with the AIA.  ‘Deployer’ defined in the Article 3(4) AIA as a natural or legal person, public authority, agency or other body using AI systems under their authority, unless the use is for a personal non-professional activity.  The AIA does not define the concept of ‘user’ which is defined in the MDR as any healthcare professional or lay person who uses a device. Therefore, the concept of ‘deployer’ under AIA cannot be understood as referring to ‘user’ under MDR/IVDR.

The MDR/IVDR requirements address risks related to medical device software, however, they do not explicitly address risks specific to AI systems. The AIA complements the MDR/IVDR by introducing requirements to address hazards and risks for health, safety and fundamental rights specific to AI systems. In line with the New Legislative Framework approach, this means a simultaneous and complementary application of the MDR/IVDR and the AIA for medical devices that contains one or more high-risk AI system.

For the purposes of this document, AI systems used for medical purposes are referred to as Medical Device Artificial Intelligence (MDAI). All references to MDAI shall be understood to also cover MDR Annex XVI products, accessories to medical devices, in vitro diagnostic medical devices and accessories to in vitro diagnostic medical devices.

The guidance provided in this document is without prejudice to the guidance that the European Commission may adopt on the basis of the AIA.

In order to ensure consistency, avoid duplication and minimise additional burdens, manufacturers of MDAI, in accordance with paragraph 2 of Article 8 of the AIA, have a choice of   integrating, as appropriate, the necessary testing and reporting processes, information and documentation they provide with regard to their MDAI into documentation and procedures that already established under the MDR/IVDR. Manufacturers of MDAI, are strongly encouraged to use this flexibility provided in paragraph 2 Article 8 of the AIA. In applying this flexibility, MDAI manufacturers, however, shall ensure that MDAI are fully compliant with all applicable requirements of the AIA, MDR or IVDR and any other applicable Union legislation.

# Scope of application and classification

- When does the AIA apply to the ‘medical device software’?

According to MDCG 2019-11, "medical device software" (MDSW) refers to software intended, either alone or in combination, to fulfil a medical purpose as defined in Article 2(1) MDR or Article 2(2) IVDR. The AI Act (AIA) defines an AI system in Article 3(1) as a machine-based system that is designed to operate with varying levels of autonomy and that may exhibit adaptiveness after deployment, and that, for explicit or implicit objectives, infers, from the input it receives, how to generate outputs such as predictions, content,  recommendations, or decisions that can influence physical or virtual environments.

Note 1: all references to Medical Device Artificial Intelligence (MDAI) shall be understood to also cover MDR Annex XVI products, accessories to medical devices, in vitro diagnostic medical devices and accessories to in vitro diagnostic medical devices.

Note 2: Refer to MDCG 2019-11 for further details on the qualification and classification of software, including the distinction between MD and IVD software and on the definition of an AI system to the Commission Guidance on the AI system definition.

- Under what conditions is a MDAI considered a high-risk AI system within the meaning of the AIA?

A MDAI is considered a high-risk AI system under Article 6(1) AIA if it meets both of the following conditions:

- the MDAI is a safety component, or the AI system is itself a medical device and

- the MDAI is subject to a third-party conformity assessment by a notified body in accordance with the MDR/IVDR (See table 1 below).

See Annex I for a qualification decision tree

| Classification | Notified Body Involved? | AIA High-Risk (Art. 6(1))? |
| --- | --- | --- |
| MDR Class I (non-sterile, no measuring, not reusable surgical) | ❌ No | ❌ No |
| MDR Class I (sterile/measuring/reusable surgical) | ✅ Yes | ✅ Yes |
| MDR Class IIa, IIb, III | ✅ Yes | ✅ Yes |
| IVDR Class A (non-sterile) | ❌ No | ❌ No |
| IVDR Class A (sterile) | ✅ Yes | ✅ Yes |
| IVDR Class B, C, D | ✅ Yes | ✅ Yes |

Table 1

- When do provisions on high-risk AI systems apply to MDR Annex XVI devices?

The AIA applies to the MDR and therefore it also applies for Annex XVI device if it is or it contains an AI system and fulfils conditions of Article 6(1) AIA (refer back to Q2).

- Does the AIA impact the risk classification of MDAI under the MDR/IVDR?

A classification of an AI system as a high-risk under Article 6(1) AIA, in accordance with recital 51 AIA, does not necessarily imply that the medical device or in vitro diagnostic medical device falls in a higher risk class under the MDR and IVDR

Therefore, the AIA does not impact the risk classification of medical devices or in vitro diagnostic medical devices (IVDs) under the MDR or IVDR.

It is rather the classification of a medical device under the MDR/IVDR which determines whether the AI system qualifies as high-risk under Article 6(1) AIA (refer back to Q2).

The classification of MDAI under the MDR or IVDR determine the applicable legal requirements for high-risk AI systems under Article 6(1) AIA and impact the regulatory scrutiny and oversight required for the MDAI under the MDR or IVDR and AIA. Both aspects influence whether requirements of the AIA apply, which AIA Article 43 conformity assessment procedure is to be followed, and whether AIA Chapter IX post-market monitoring requirements apply. Both aspects also impact AIA obligations for manufacturers in accordance with Article 16, and other economic operators, such as deployers and distributors.

The application of Article 5 AIA, prohibited AI practices and Article 50 AIA transparency obligations for providers and deployers of certain AI systems does not depend on the MDR/IVDR classification.

- How does the AIA differentiate AI systems that are or that are part of MDAI?

In the classification rule for high-risk AI systems, the AIA differentiates between AI systems that are intended to be used as a safety component or are itself a product. A safety component in defined in Article 3(14) AIA. European Commission is preparing horizontal guidelines on the classification of AI systems as high risk that will also cover the concept of ‘safety component’. .

# Requirements

## Management Systems

- How do the MDR/IVDR and AIA address the lifecycle management for MDAI?

The MDR and IVDR require manufacturers to manage the entire lifecycle of MDAI, ensuring that the MDAI remains safe and performant throughout its use (MDR/IVDR, Annex I, Chapter II)

The AIA reinforces this by expecting continuous review, oversight, and consistent performance of MDAI throughout the lifecycle (Article 9 para 2, Article 15, par. 1 AIA) and post-market monitoring that is proportionate to the nature and risks of the AI system (Article 72, par. 1, AIA). This includes design choices allowing  natural persons to oversee functioning of AI systems and to ensure that they are used as intended and that their impacts are addressed over the system’s lifecycle (Article 14 and recital 73 AIA), risk management system that shall be understood as a continuous iterative process planned and run through the entire lifecycle of a high-risk AI system, requiring regular systematic review and updating,  analysing data provided by deployers or collected through other sources on the performance of the MDAI in its lifetime to assess continued compliance. This may include activities such as MDAI design, development, testing, deployment, monitoring, and updates, along with comprehensive documentation of all relevant changes and their impacts.

For MDAI that continues to learn after being deployed, the post-market monitoring system is key to ensuring continued performance and compliance. It helps identify and address emerging  risks in a timely and efficient manner and ensures that updates are communicated  to relevant parties, including patients or professional users( as applicable) , deployers and notified bodies, as necessary. (AIA Recital 155, Article 15(4)). Article 9, 11, 15, Annex IV.

- How do the MDR, IVDR and the AIA address the implementation of a quality management system?

Jointly, the AIA and MDR or IVDR emphasize the importance of quality management to ensure the safety and performance of MDAI. The MDR and IVDR requires manufacturers to establish, document, implement, maintain, keep up to date and continually improve a quality management system to ensure compliance with the MDR and IVDR in the most effective manner and in a manner that is proportionate to the risk class and the type of the device.

Similarly, the AIA mandates in Article 17 that providers/ manufacturers implement quality management systems for AI systems to ensure compliance with the AIA. The quality management system required under the AIA covers substantive requirements and procedural obligations related to the quality of the AI system. It includes a list of at least thirteen aspects, including a strategy for regulatory compliance, the documentation of a quality management system, which should encompass written policies, procedures, and instructions regarding aspects such as risk management and performance testing (AIA, Article 17).

The quality management system obligations under the AIA are specifically targeted to the AI system. Therefore, it is complementary to the quality management system required under the MDR or IVDR and applicable to the overall MDAI.

Accordingly, to ensure this complementarity and to avoid unnecessary administrative burden, Recital 81, Article 8 (2) and Article 17 (3) AIA clarify that manufacturers of AI systems may include the elements of the quality management system provided by the AIA as part of the existing quality management system provided by the MDR and IVDR. Regardless of the approach employed, manufacturers of MDAI must ensure that Article 17 requirements of the AIA, including compliance of MDAI with the requirements set in Section 2 Chapter III, such as risk-management system, data and data governance, technical documentation, record-keeping, transparency, human oversight, accuracy, robustness and cybersecurity, are fully covered in the quality management system provided by the MDR and IVDR.

In addition, the development of harmonised European and international standards relevant to quality management systems for MDAI is ongoing, including under the framework of the AIA and MDR/IVDR, to support consistent implementation and compliance across manufacturers.

- What risk management requirements are outlined in the MDR, IVDR and AIA for MDAI?

The MDR and IVDR require a risk management system that shall be understood as a continuous iterative process throughout the entire lifecycle of the device, including pre-market and post-market phases (Annex I Chapter I MDR, IVDR).  The process is aiming primarily at identifying and mitigating risks on health and safety.

The AIA complements the risk management process required by the MDR/IVDR and introduces AI system specific risk management requirements (Article 9 AIA). Similar to the MDR/IVDR, the AIA requires the risk management system for AI system to be a continuous, iterative process that is planned and run through the entire lifecycle of a high-risk AI system. That process should be aimed at identifying and mitigating relevant risks of AI systems on health, safety and fundamental rights.

Additional risk management requirements specific to the AI system include ongoing assessments of known and reasonably foreseeable risks that the high-risk AI system can pose to health, safety and fundamental rights, data biases, and system robustness, including identification, analysis, and mitigation of risks related to system design, development, and deployment, and may include training to deployers. This includes measures to ensure the safety and reliability of AI systems in healthcare settings through comprehensive risk assessments, documentation of risk mitigation measures, and continuous monitoring of system performance (AIA, Article 72). These requirements aim to prevent, and address known and reasonably foreseeable risks to ensure the safety and performance of MDAIs.

Manufacturers of MDAI may integrate the additional risk management requirements specific to an AI system set out in paragraphs 1-9 of Article 9 of the AIA into the  testing, reporting processes, information and documentation required under the AIA into their existing documentation and procedures required under the MDR and IVDR (recital 64 AIA). Risk management systems under both AIA and MDR/IVDR require a manufacturer to reduce the identified and assessed risks, and to address risks in their risk management systems that can reasonably be mitigated.  This refers not only to organisational measures, but also to specific actions taken during the development and design of AI systems.

## Data Governance

- What data governance requirements are specified in the AIA, MDR and IVDR for MDAI?

The MDR/IVDR mandates that clinical data used for device evaluation is robust, reliable, and derived from well-designed studies. To support the generation of sufficient clinical evidence, the clinical evaluation / performance evaluation must be based on clinical data representative of the intended use of the device. The procedures and techniques for verifying, validating and controlling the design of the devices as well as documented information arising from those procedures and techniques, should be documented as part of the clinical evidence (and reviewed by the notified body).

Article 10 AIA provides further specifications related to the data and datasets of AI systems. Moreover in Article 10(5) AIA includes a  provision for exceptional procession of special categories of personal data. These requirements aim to ensure that MDAIs are built and validated on trustworthy data.

To ensure that MDAIs performs as intended and safely, high-quality data plays a vital role in providing structure and in ensuring their performance, especially when for the development of MDAI the the training of AI models are used. High-quality data sets for training, validation and testing of AI-systems require the implementation of appropriate data governance and management practices. The data governance and management practices, in the case of personal data, should be fully compliant with the GDPR provisions and include transparency about the original purpose of the data collection. The requirements related to data governance can be complied with by having recourse to third parties that offer certified compliance services including verification of data governance (Recital 67 AIA).

The AIA further mandates comprehensive data governance practices for MDAIs, specifying that datasets for training, validation, and testing must be relevant, sufficiently representative, and, to the best extent possible, free of errors and complete (AIA Article 10, recital 67). When training, validating and testing the MDAIs on datasets, manufacturers/ must ensure that the datasets are sufficiently representative of the target population, and specific the selection criteria employed as well as the decisions related to the size of the performance study population (see IVDR Annex XIII part 2.3.2 (m); Article 10 para 3 AIA).  Manufacturers must implement procedures ensuring data transparency and integrity and examine the data in view of possible biases, with detailed documentation of compliance.

For further information on generation on clinical evaluation (MDR) / performance evaluation (IVDR) of MDSW, see MDCG 2020-1.

- What requirements are in place under the AIA, MDR and IVDR to ensure monitoring and mitigation of unwanted bias in MDAI?

The AIA requires MDAI manufacturers to implement data governance and management practices appropriate for the claimed intended purpose, including with a view of possible biases that are likely to affect the health and safety of persons, have a negative impact on fundamental rights or lead to discrimination prohibited under Union law, especially where data outputs influence inputs for future operations

Furthermore, manufacturers must implement appropriate measures to detect, prevent and mitigate possible biases identified (Article 10 (g) and (f) AIA).

In addition, Article 12 AIA outlines requirements on record keeping arising from logging capabilities implemented for MDAI that can facilitate the traceability i.e. identification of situations whereby an MDAI may present a risk due to potential bias in the training, validation or testing data sets in the initial system development or as a result of substantial modification. The AIA, as one of the essential requirements, requires all high-risk AI systems have technical capabilities for the automatic recording of events (logs) over the lifetime of the MDAI.

The record-keeping and logging requirements under the AIA also involve sufficiently representative datasets analysing relevant data provided by the deployers (Article 72, par. 2 AIA). The MDR and IVDR complements this by requiring that clinical data used in device evaluation be robust and reliable, ensuring that MDAI perform consistently across the intended use population (MDR, IVDR Annex XIV). Documentation (e.g. Annex II and III) of these activities and their effectiveness is required by all three regulations (see also next questions).

Furthermore, the requirements for appropriate record keeping via logging capabilities as part of post-market monitoring (as referred to in Article 72 AIA and MDR/IVDR PMS requirements + MDR/IVDR Annex III) aim to ensure that manufacturer put in place appropriate mechanisms to detect possible bias not originally detected in pre-market activities but rather as a result post-market monitoring or learning.

- How does the AIA define different types of data required to demonstrate compliance of the AI system?

Article 3 (29) AIA defines ‘training data’ as data used for training an AI system through fitting its learnable parameters;

Article 3 (30) AIA ‘validation data’ means data used for providing an evaluation of the trained AI system and for tuning its non-learnable parameters and its learning process in order, inter alia, to prevent underfitting or overfitting;

Article 3 (31) AIA ‘validation data set’ means a separate data set or part of the training data set, either as a fixed or variable split;

Article 3 (32) AIA ‘testing data’ means data used for providing an independent evaluation of the AI system in order to confirm the expected performance of that system before its placing on the market or putting into service.

- How do the MDR, IVDR and AIA address training data used for MDAI?

The importance of utilising training data which is appropriate for the intended purpose of the MDAI is essential in order to produce accurate and (clinically) relevant outputs. In accordance with the claimed intended purpose of the device, the training data used must be representative of the intended patient population. (MDR Annex II TD section 3, a and b and section 6 on product verification and validation).

The AIA requires that training data for MDAI be of high quality, sufficiently representative, free of errors (to the best extent possible), complete in the view of the intended purpose , include the appropriate statistical properties and be examined in view of possible biases that are likely to affect the health and safety of persons, have negative impact on fundamental rights ore lead to discrimination prohibited under Union law  (AIA, Article 10). It also requires the inclusion of measures to address data privacy and security concerns, as well as transparency in data collection and processing.

The MDR and IVDR require that clinical data used for device evaluation is robust, reliable, and derived from well-designed studies (MDR, Art. 57 (3)).). Annex XIV - MDCG 2020-1 Guidance on clinical evaluation (MDR) / Performance evaluation (IVDR) of medical device software).

Should the device have a broad intended purpose, then a diverse, representative dataset, free from biases to the extent possible should be utilised to generate sufficient clinical evidence. Particular importance should be placed on the diversification of datasets that should cover the various demographics and conditions claimed.

Manufacturers should employ stringent data governance practices to maintain data integrity and prevent unwanted bias (AIA, Article 10). In addition, the validation of training data used MDAI is paramount and should be demonstrated as part of the studies to ensure the accuracy, reliability, and effectiveness of the MDAI.

## Technical Documentation

- What technical documentation is required by the MDR/IVDR and AIA for MDAI?

The MDR, IVDR and the AIA mandate the provision of comprehensive technical documentation for MDAI. The MDR/IVDR require detailed descriptions of software, software architecture, data processing methods, and risk management strategies (MDR, Annex II). The AIA requires additional documentation, focusing on transparency and accountability, including risk assessments, data governance practices, and performance testing outcomes of the AI system (AIA, Article 11).

This documentation should include detailed information about the device's design, development, key design choices, functionality, performance characteristics, system architecture, computational resources to develop, train and test, and intended use and purpose. Manufacturers must also provide evidence of conformity with relevant regulatory requirements, including training, validation and test data, risk assessments, and quality management processes (MDR, Annex II; AIA, Article 11). Both regulations aim to ensure that manufacturers maintain detailed and up-to-date records to demonstrate compliance. The technical documentation should also address the requirements of MDR Annex III and Article 72 AIA.

In accordance with Article 8 paragraph2 of the AIA to avoid duplication and minimize additional burdens manufacturers have a choice of integrating, as appropriate, the necessary testing and reporting processes, information and documentation required under the AIA with regard to their MDAI product into documentation and procedures that already exist and are required under the MDR or IVDR.

- Will the assessment of the technical documentation by a notified body as it is defined in Annex VII of the AIA follow a sampling of MDR Class IIa / Class IIb and IVDR Class B / Class C devices as it is laid out in MDCG 2019-13?

Yes, sampling rules of the governing conformity assessment procedure remain applicable. AI systems falling under the definition of an in vitro diagnostic or medical devices of the IVDR or MDR are governed by the applicable regulation. Since Article 52(4) and (6) of the MDR and Article 48(7) and (9) of the IVDR establish the need to assess the technical documentation of at least one representative device per generic device group (for Class IIb and Class C) and for each category of devices (for Class IIa and Class B) prior to issuing a certificate. Section 2.3 and 3.4 of Annex IX of IVDR and MDR (and section 10 of Annex XI of the MDR) defines that the quality management system assessment must be accompanied by the assessment of technical documentation for devices selected on a representative basis.

## Transparency and human oversight

- What requirements do the AIA and MDR/IVDR impose regarding transparency and explainability for MDAI?

The AIA and MDR/IVDR put in place complementary obligations that address both manufacturers and deployers to ensure transparency in the development, deployment and use of MDAI systems.

Under the AIA, transparency is a core requirement for high-risk AI systems. Article 13 AIA introduces on providers a legally binding requirement to design and develop MDAI in such a way as to ensure that operation of MDAI is sufficiently transparent to enable deployers, to interpret outputs correctly and use the system appropriately, supported by clear and comprehensible instructions for use. Additionally, Article 26 imposes transparency obligations on deployers, including the requirement to inform providers appropriately and ensure proper use of the system.

The MDR/IVDR embed transparency requirements within the General Safety and Performance Requirements (GSPRs), particularly in Annex I. GSPR 23 requires that manufacturers provide clear and accessible information regarding the device’s intended purpose, operation and limitations. GSPR 14.2(h) requires that software development follows the state of the art, incorporating lifecycle and risk management processes that inherently support traceability, documentation, and usability. These elements contribute to the transparency and the broader goal of ensuring that deployers and regulators can understand how the AI component of a device contributes to its performance and risk profile.

Annex II and III of the MDR/IVDR also impose detailed documentation requirements, including on software development and performance evaluation, which contribute to transparency and traceability throughout the device lifecycle. Accordingly, the MDR/IVDR and AIA collectively support a coherent regulatory framework that ensures MDAI systems are designed, documented, and deployed in a transparent and explainable manner.

The MDR/IVDR require manufacturers to ensure that users are provided with comprehensive and comprehensible information regarding device performance, risks, and clinical evidence (Annex I). This includes transparent reporting of risk assessments, performance testing, and post-market surveillance data. Where applicable, the information provided must describe how software, including AI components, contributes to the performance of the device, and must be reflected in the instructions for use or user interface

Therefore, transparency requirements are not optional design features but essential requirements to be addressed within the manufacturer’s risk and quality management systems and verified through the conformity assessment procedure. These requirements collectively ensure that users, deployers, and patients are adequately informed about the nature, operation and limitations of MDAI systems.

- How do the MDR/IVDR and the AIA address transparency, explainability and data processing requirements for MDAI?

Articles 12 and 13 AIA as clarified in recitals 71 and 72, provide that transparency, including the accompanying instructions for use, should assist deployers in the use of the system and support informed decision making by them. Deployers should, inter alia, be in a better position to make the correct choice of the system that they intend to use in light of the obligations applicable to them, be educated about the intended and precluded uses, and use the AI system correctly and as appropriate. This is further supported by requirements for clear and comprehensible instructions for use, the provision of information on the system’s capabilities and limitations, and documentation that enables the explainability of AI-based decisions.

Article 10 AIA sets transparency requirements related to data processing. This article includes obligations regarding data quality, management, and documentation. Data used for training, validation and testing of the AI system must be relevant, representative of the intend purpose, free of errors and bias, (to the extent possible) and sufficiently comprehensive. These provisions aim to ensure not only the robustness and performance of MDAI but also their transparency in terms of data processing.

This is aligned with MDR/IVDR which require manufacturers to ensure that users are provided with comprehensive and comprehensible information regarding the device, its performance and risks (Annex I). This includes transparent reporting of risk assessments, performance testing, and post-market surveillance data. Where applicable, the information provided must describe how software, including AI components, contributes to the performance of the device, and must be reflected in the instructions for use or user interface. Furthermore, MDR Annex I,

Chapter III, also supports informed decision-making by users.

In this sense, the AIA expands and reinforces the MDR’s foundational principle that information provided to deployers must be clear, complete, and actionable, especially when decisions may impact health outcomes or fundamental rights. Taken together, the MDR and AIA establish a framework in which transparency and explainability are not only expectations but binding obligations for MDAI. Manufacturers must ensure that regulators and deployers can understand the logic, limitations, and behaviour of AI components within a medical device. This includes implementing safeguards, interpretation tools, and user interfaces that make AI outputs meaningful and trustworthy. The integration of these requirements throughout the MDAI lifecycle, design, documentation and record-keeping, labelling, and post-market surveillance ensures that MDAI is developed and deployed in a manner that supports patient safety, professional accountability, and public trust.

- How do the MDR/IVDR and the AIA address the accountability of MDAI?

Accountability is addressed in both the AIA and the MDR/IVDR, though with a different scope and focus.

The MDR/IVDR, includes requirements for documentation and clinical and performance evaluation. Annex I requires that the information provided to users be clear and complete, including details on how the device operates and, where applicable, how embedded software components (including AI) contribute to device functionality. The documentation required under Annex II must describe the design and functioning of software, including how inputs are processed and how outputs are generated, thereby supporting explainability and accountability in practice. To promote trust and accountability, the AIA introduces explicit obligations concerning transparency in Article 13. Transparency contributes to explainability which in turn facilitates accountability. As such, high-risk MDAI must be designed and developed in a manner that enables deployers to understand how the system functions and reaches its outputs. This includes information on the characteristics, capabilities, and limitations of the MDAI, as well as documentation to support interpretability of outputs.

These combined obligations enable both developers and deployers of MDAI to demonstrate and communicate how MDAI-based decisions are made. Moreover, these obligations support robust traceability, informed use, and post-market control mechanisms that reinforce the safe and trustworthy deployment of MDAI and enhance accountability throughout the product lifecycle.

- How do the MDR, IVDR and AIA address usability engineering for MDAI?

The MDR and IVDR require manufacturers to apply usability engineering principles to MDAI in the design and development to ensure their safe and effective use by intended users. Manufacturers must eliminate or reduce as far as possible risks related to use errors, having particular regard to the knowledge of the user and consideration of whether training might be appropriate. It is recommended that AI systems, especially those in healthcare, be designed with user-centric principles to facilitate safe and effective interaction. Documentation of usability engineering processes and outcomes is required under all three regulations.

- What human oversight requirements are included in the MDR, IVDR and AIA for MDAI?

The AIA emphasizes the importance, and sets a legal obligation to design and develop AI systems, including MDAI, with appropriate  human oversight mechanisms. In particular human oversight measures should guarantee that the system is subject to in-built operational constraints that cannot be overridden by the system itself and is responsive to the human operator, (AIA, Article 14). This involves designing MDAI to allow human intervention in critical decision-making processes. The MDR supports this by requiring that medical devices, including MDAI, be designed for safe and performant use (MDR, Annex I). The AIA details that oversight measures should be commensurate with the risks, level of autonomy, and context of use of the MDAI. Thus, clearly defined and documented human oversight mechanisms as well as appropriate instructions for use (MDR, Annex I) will be necessary to ensure safe and performant use and deployment of MDAI and allow appropriate supervision by healthcare professionals and institutions.

- Can human oversight for medical devices be understood as a part of existing risk management measures such as a design measure (e. g. ‘stop’ button)?

Human oversight in additional to the design and operational requirement (see Q17), is also considered a risk mitigating factor that aims to prevent or minimise the risks to health, safety or fundamental rights when a high-risk AI system is used in accordance with its intended purpose or under conditions of reasonably foreseeable misuse according to Article 14(2) of the AIA. Proper human oversight could be understood as a risk management measure according to Annex I (4) MDR/IVDR which calls for manufacturers to:

- eliminate or reduce risks as far as possible through safe design and manufacture take adequate protection measures, and

- (where appropriate) include including alarms, if necessary, in relation to risks that cannot be eliminated and

- provide information for safety.

For medical devices, specific considerations as to what level of human oversight is necessary and appropriate according to the level of risk associated with possibly e.g. robotic driven surgical medical device intervention where a healthcare professional oversees the operation and has variable possibilities to ‘override’ the MDAI. However, the software design should not allow human intervention in critical parts of the surgical operation whereby leaving the patient at risk. The manufacturer must include considerations as part of the risk assessment and management.

.

- How do the MDR/IVDR and AIA approach informed consent in the context of MDAI?

The MDR/IVDR and AIA both include provisions aimed at protecting individual rights through transparency and informed interaction with MDAI.

Under the MDR and IVDR, informed consent is explicitly required in the context of clinical investigations and performance studies (MDR Article 69, IVDR Article 59). These provisions ensure that individuals participating in such studies are adequately informed of the risks, benefits, and objectives of MDAI.

The AIA complements these requirements by introducing additional transparency obligations that extend to the general deployment of AI systems. Article 50(1) requires that individuals interacting with an AI system be informed that they are doing so. Furthermore, Article 14 on human oversight and Article 13 on transparency reinforce the need to provide deployers and affected persons with sufficient information to understand the system’s capabilities, limitations, and potential risks.

Together, these requirements contribute to safeguarding patient autonomy and support the ethical deployment of MDAI systems.

- How do the MDR/IVDR and AIA address the traceability of AI algorithms used in MDAI?

Traceability is a key component of both regulatory frameworks, albeit with different focal points.

The MDR (Article 25) and IVDR (Article 22(1)) require that devices, including those incorporating AI, are traceable throughout the supply chain and device lifecycle. This includes obligations related to Unique Device Identification (UDI), registration, and post-market surveillance.

With the same intent, under the AI Act the EU declaration of conformity, among other information, requires “AI system name and type and any additional unambiguous reference allowing the identification and traceability of the AI system” to be included. Moreover, Article 12 AIA

introduces requirements related to functional traceability. Article 12 mandates that high-risk AI systems maintain logs of system performance and behaviour throughout their lifecycle to support monitoring and post-market monitoring. Article 12 and recital 71 underscore that logging and documentation are essential to ensure system traceability of the functioning of a high-risk AI system.

Thus, the concept of traceability is applied in two interrelated ways: (1) traceability of device movement and lifecycle and (2) traceability of system functioning and performance. Together, these ensure that both the hardware and software dimensions of MDAI are adequately monitored and controlled.

## Accuracy, Robustness and Cybersecurity

- What cybersecurity measures are required by the AIA and MDR/IVDR?

The MDR ,IVDR and AIA emphasize the need for robust cybersecurity measures in both the pre-market and post-market stages of MDAI. The MDR and IVDR underline that any risks associated with the operation of the device must be acceptable so as to enable a high level of protection of health and safety, taking into account the generally acknowledged state of the art. This can only be achieved through the establishment of an adequate balance between benefit and risk during all possible operation modes of a device. To this end, there is a need to consider the relationship between "safety” and “security" related risk. All three regulations require manufacturers to establish procedures to consider safety and security risk aspects from an early stage of design and throughout the entire life cycle, taking into account the operational environment of use and ICT infrastructure. (Based on MDCG 2019-16 rev. 1 and Recital 76 AIA).

The cybersecurity measures implemented by manufacturers must aim to prevent unauthorised access, cyberattacks, exploits, manipulation and ensure operational resilience. Article 15 AIA requires for high-risk AI systems technical solutions to address AI specific vulnerabilities. Article 15 is part of the essential requirements for high-risk AI systems, and thus compliance of MDAI with cybersecurity requirements is an obligation of the manufacturer in accordance with Article 9, it should be part of the risk management system, in accordance with Article 16 (a), it should be part of the quality management system in accordance with Article 17 and be subject to conformity assessment in accordance with Article 43. As further explained in recitals 76 and 77 of the AIA, manufacturers of high-risk AI systems should ensure a level of cybersecurity appropriate to the risks and take suitable measures to secure AI specific assets such as training data sets or trained model as well as appropriate the underlying ICT infrastructure.

Manufacturers of MDAI should implement measures to secure data transmission and storage, prevent unauthorised access, data and model poisoning, and detect and respond to cybersecurity incidents. Manufacturers must also conduct risk assessments to identify potential cybersecurity vulnerabilities and implement appropriate mitigation measures (MDR, Annex I, Section 17; AIA, Article 15, Annex IV, Section 2.h). The technical solutions aiming to ensure the cybersecurity of MDAIs shall be appropriate to the relevant circumstances and the risks.

The technical solutions aiming to ensure the cybersecurity of high-risk AI systems shall be appropriate to the relevant circumstances and the risks.

The MDR and IVDR underline that any risks associated with the operation of the device must be acceptable so as to enable a high level of protection of health and safety, taking into account the generally acknowledged state of the art. This can only be achieved through the establishment of an adequate balance between benefit and risk during all possible operation modes of a device. To this end, there is a need to consider the relationship between "safety” and “security" related risk. The AIA, MDR and IVDR require manufacturers to establish procedures to consider safety and security risk aspects from an early stage of design and throughout the entire life cycle, taking into account the operational environment of use and ICT infrastructure. (Based on MDCG 2019-16 rev. 1 and Recital 76 AIA).

In addition, requirements regarding privacy and confidentiality of data associated with the use of MDAIs that may be outside the scope of the MDR and IVDR but are subject to other legislations, such as NIS-2 Directive[1] and GDPR[2].

Note:  medical devices and in vitro diagnostic medical devices are out the scope of Regulation (EU) 2024/2847 of the European Parliament and of the Council of 23 October 2024 on horizontal cybersecurity requirements for products with digital elements and amending Regulations (EU) No 168/2013 and (EU) 2019/1020 and Directive (EU) 2020/1828 (Cyber Resilience Act).

[1] Directive (EU) 2022/2555 of the European Parliament and of the Council of 14 December 2022 on measures for a high common level of cybersecurity across the Union, amending Regulation (EU) No 910/2014 and Directive (EU) 2018/1972, and repealing Directive (EU) 2016/1148

[2] Regulation (EU) 2016/679 of the European Parliament and of the Council of 27 April 2016 on the protection of natural persons with regard to the processing of personal data and on the free movement of such data, and repealing Directive 95/46/EC

# Clinical /Performance Evaluation and Testing

- What criteria is specified in the AIA for evaluating the performance of MDAI?

The AIA specifies criteria for ensuring the safety, reliability, and effectiveness of MDAI. While the AIA does not explicitly use the term “clinical evaluation” or "performance evaluation," it mandates requirements such as accuracy, robustness, and cybersecurity for high-risk AI systems, which are essential aspects of performance (MDR, Article 61, Annex XIV, IVDR Article 56; AI Act, Article 15, recital 74) as well as specifically requires testing of high-risk AI systems against prior defined metrics and probabilistic thresholds to ensure that high-risk AI systems perform consistently for their intended purpose and that they are in compliance with the requirements of the AIA (Article 9).

The AIA mandate validations of AI training pipelines to ensure the reliability and accuracy of the AI models. This includes validating design, manufacturing, data collection, preprocessing, model training, and quality management processes under various conditions, with continuous monitoring to address potential issues (AIA, Article 10 and 13).

- What specific requirements do the MDR, IVDR and AIA impose for the clinical (MDR) or performance (IVDR) evaluation of MDAI?

The MDR and IVDR require manufacturers to validate MDAI outputs through rigorous testing in the form of clinical or performance evaluation to ensure compliance and that the MDAI performs as intended (MDR Article 61, IVDR Article 56, Annex XIV). Manufacturers must perform software verification and validation activities to ensure that MDAI meets specified requirements and functions correctly. This requires clinical validation to demonstrate that the MDAI is safe and provides accurate, reliable, and clinically relevant outputs (MDR IVDR Annex XIV).  Similarly, the AIA mandates verification and validation of MDAI to ensure it operates as intended and meets safety and performance requirements (See Q24). In addition to the MDR and IVDR, the AIA introduces the requirement for the validation of high-risk MDAI, in terms of transparency, human oversight, accuracy, robustness and cybersecurity (Chapter III, Section 2) . Furthermore, the AIA requires that validation of MDAI involves verifications that the MDAI does not infringe on fundamental rights.

All three frameworks emphasize testing under various conditions, documentation of evaluation processes, and continuous monitoring to ensure compliance. These regulations require manufacturers to provide evidence to support their claims, ensuring that MDAI adhere to rigorous safety and performance standards.

This includes testing, documentation of validation processes, and continuous monitoring to address any issues that may arise (AIA, Articles 9, 10, 12 and 17). Both regulations provide flexibility in choosing validation methods but emphasize the importance of demonstrating that the MDAI functions correctly and meets specified requirements.

Refer to MDCG 2020-01 “Guidance on Clinical Evaluation (MDR) / Performance Evaluation (IVDR) of Medical Device Software”.

- What processes are outlined in the MDR, IVDR and AIA for generating clinical evidence to support the safety and performance of MDAI?

The MDR and IVDR introduce requirements with regard to the generation of clinical evidence to support the safety and performance of MDAI. The MDR /IVDR require manufacturers to generate clinical evidence through the conduct of clinical/performance evaluations, which may be generated through clinical investigations and clinical performance studies (as applicable) to demonstrate conformity with the applicable requirements (MDR, IVDR Annex XIV). Under the AIA, real world testing may be necessary to assess the safety and effectiveness of the MDAI. This involves designing and conducting studies to evaluate the performance, reliability, and (clinical) impact of MDAI, with specific requirements for study design, data collection, and statistical analysis. Both regulations emphasize the importance of generating robust clinical evidence to demonstrate the safety and effectiveness of MDAI.

- What requirements are outlined in the MDR and AIA for the technical/ clinical evaluation of AI systems used in medical devices?

The MDR requires manufacturers to provide robust clinical data to support claims about the device’s performance and safety, ensuring compliance with MDR requirements for clinical evaluation (MDR, Annex XIV) and clinical evidence. The MDR / IVDR outline specific requirements for the clinical evaluation of AI systems used in medical devices to ensure their safety, effectiveness, and clinical benefit. It also requires (where applicable) manufacturers to conduct clinical investigations and performance evaluations to assess the safety and performance of medical devices, including AI-enabled ones (MDR, Annex XIV). Similarly, the AIA mandates to include in the quality management system techniques, procedures and systemic actions to be used for the design, design control and design verification as well as techniques, procedures and systemic actions to be used for the development, quality control and quality assurance of the high-risk AI systems (Article 17) with the specific requirements related to data (Article 10, Article 17). Article 43 AIA introduces additional consideration to be taken into account as part of the clinical or performance evaluation, for high-risk AI systems that continue to learn after being place on the market or put into service, such as pre-determined changes to the MDAI.

- How might manufacturers conduct clinical investigations and clinical performance studies in conformity with both MDR / IVDR and the AIA?

High-risk

MDAI must be supported by clinical evidence to demonstrate the device’s safety, performance, and, where applicable, clinical benefit. In accordance with the MDR and IVDR, such clinical evidence may need to be generated through a clinical investigation (under the MDR) or a performance study (under the IVDR), as appropriate.

Where an MDAI undergoes a clinical investigation or performance study, this constitutes real-world testing as referred to in Article 60(1) of the AIA. In such cases, paragraph 3 of Article 60(1) applies, which allows real-world testing of high-risk AI systems prior to their placing on the market or putting into service, provided that such testing is in accordance with applicable Union or national laws, including those under the MDR and IVDR.

Although the AIA does not apply to AI systems that are not yet placed on the market or put into service (Article 2(8) AIA), Article 60 of the AIA introduces a specific provision for Annex III high-risk AI systems, such as MDAI. Under certain conditions, these systems may undergo testing in real-world conditions prior to their placing on the market or putting into service. Article 60(1), third subparagraph, explicitly states that such real-world testing is permitted "without prejudice to Union or national law on the testing in real-world conditions of high-risk AI systems that are medical devices or in vitro diagnostic medical devices.”

# Conformity assessment

- Which conformity assessment procedure apply to AI systems in the scope of the MDR/IVDR and AIA?

In accordance with Article 43 (3) AIA, for high-risk MDAI (classified under Article 6(1)) the relevant conformity assessment procedure is determined by the MDR/IVDR.

An AI system that is classified as high-risk based on Article 6(2) AIA, and falls into one of the areas listed in Annex III points 2 to 9, in accordance with paragraph 2 Article 43 (2) AIA, shall follow the conformity assessment procedure set in Annex VI AIA. This is the procedure based on internal control which does not provide for the involvement of a notified body.

An AI system that is classified as a high-risk based on Article 6(2) AIA, and falls into area listed in Annex III point 1, in accordance with Article 43(1) AIA, shall follow one of the conformity assessment procedures provided in this article.

For example, if an AI system used for healthcare, such as an emergency triage system, qualifies as a medical device under MDR, then conformity assessment procedure following AIA Art. 6 (1) is applicable. AI systems that qualify as high-risk AI systems under Article 6(1) AIA must comply with both MDR or IVDR and the AIA.

However, if an AI system qualifies solely as a biometric categorisation system under Annex III Section 1 (b) AIA, then this system must comply only with the requirements and obligations of the AIA and undergo a conformity assessment procedure in accordance with Article 43 (1) AIA.

- What is the process for demonstrating conformity to both the AIA and MDR/IVDR?

Following Article 16(f) and Article 43(3) AIA, manufacturers of high-risk MDAI undergoes the  relevant conformity assessment procedure based on the device's risk classification as described in Article 52 MDR or Article 48 IVDR. Most MDAI are classified as Class IIa (MDR), B (IVDR), or above, meaning that they require a notified body to conduct a quality management system audit, technical documentation review, and inspections to ensure compliance with all regulatory requirements.

According to Article 43(3) AIA, high-risk MDAI the requirements set in Articles 8 to 15 AIA and specific provisions related to the assessment of quality management system and technical documentation covered in Annex VII point 4.3, 4.4, 4.5, and the fifth paragraph of point 4.6 AIA must be assessed or taken into consideration as part of the conformity assessment procedure under MDR and IVDR.

# Substantial Modification/significant change

- Will the type of substantial modifications under AIA be aligned with the changes that could necessitate a new conformity assessment under MDR/IVDR? What is the impact of substantial modifications under AIA to an AI system registration under MDR/IVDR?

. The concept of substantial modification is an autonomous concept under the AIA which is defined in Article 3 (23). In accordance with Article 43(4) of the AIA, MDAIs that have been subject to a conformity assessment procedure shall undergo a new conformity assessment procedure in the event of a substantial modification, regardless of whether the modified system is intended to be further distributed or continues to be used by the current deployer. In accordance with Article 96 (1) (c) AIA, the Commission will develop guidelines on the practical implementation of the provisions related to substantial modification.

- When do changes to a MDAI after being placed on the market or put into service not constitute a substantial modification under AIA and MDR/IVDR?

In addition to MDR/IVDR requirements for the implementation of procedures to manage updates and modifications to MDAIs, MDAIs that continue to learn after being placed on the market or put into service have the possibility of having predetermined changes plan checked at the time of the initial conformity assessment (Article 43(4) AIA).  Changes to MDAIs that have been pre-determined by the manufacturer and assessed at the moment of the initial conformity assessment and are part of the information contained in the technical documentation referred to in AIA Annex IV point 2(f) shall therefore not constitute a substantial modification.

Therefore, such pre-determined change should not be understood as a change to the certified medical device or IVD under MDR Annex IX Section 4.10 and IVDR Annex IX Section 4.11. The above-specified rule should be included in the change management procedure of the MDR/IVDR device and in the technical documentation as provided in AIA Annex IV point 2(f). It is essential that at the time of conformity assessment, the predetermined changes must be clearly specified and adaptable to the device's evolving nature. This performance, documented in the technical documentation, should include a detailed description of the initial performance expectations alongside mechanisms for validating and managing changes that occur post-market.

Note: Guidance on Pre-determined Change Control Plans for MDAI is currently under development at the level of the International Medical Device Regulators Forum (IMDRF) which will likely serve as a basis for future EU guidance on the subject-matter.

- Do high-risk MDAI that are already on the market and that undergo a significant change in the design (Article 111 AIA) before 2 August 2027 need to undergo a new conformity assessment following the entry into force of the AIA?

No.

Considering that as per the ‘Blue Guide’ on the implementation of EU product rules 2022[1], the term “placing on the market” refers to each individual product rather than to a type of product.

AIA Art. 6(1) sets out the conditions for an “Annex I high-risk AI system”. The date of application of obligations for “Annex I high-risk AI systems” is 2 August 2027, meaning:

- If the medical device has been placed on the market/put into service before 2 August 2027:

  - If the AI system is subject to any significant changes in its design on or after 2 August 2027, EU AIA obligations apply, including obligations for “Annex I high-risk AI systems”

  - If the AI system is subject to any significant changes in its design before 2 August 2027: obligations for “Annex I high-risk AI systems” do not yet apply

- If the medical device is placed on the market/put into service on or after 2 August 2027, then EU AIA obligations apply, including obligations for “Annex I high-risk AI systems”

Note: Article 111 (2) AIA gives provisions for high-risk AI systems already placed on the market or put into service and says that the EU AIA applies to high-risk AI systems that are subject to significant design changes already as of 2 August 2026. However, because EU AIA, Art. 6 (1) only applies as of 2 August 2027 (as specified in EU AIA, Art. 113 (c)), AI systems that are part of a medical device (or that are themselves medical devices) are not classified as high-risk AI systems prior to 2 August 2027, and therefore Art. 111 (2) does not apply to them prior to 2 August 2027. It is, however, to be assumed that Article 111 (2) applies to such “Annex I high-risk AI systems” by analogy, i.e. for significant changes on or after 2 August 2027.

# Post-market monitoring

- What post-market surveillance requirements are outlined in the MDR and AIA for MDAI?

Both the MDR and AIA mandate that manufacturers establish and implement post-market monitoring and surveillance systems to monitor the performance and safety of MDAI after they are placed on the market. This includes systematic collecting and analysing data on device performance, risk analysis, adverse events assessment and reporting, and other safety-related issues, as well as taking appropriate corrective and preventive actions as necessary. Manufacturers must also implement a vigilance system, report adverse events and other safety-related information to regulatory authorities and users as required (Article 83 MDR; Article 72 AIA ). Manufacturers must regularly update their risk and quality management systems and compliance strategies based on post-market monitoring and regulatory feedback.

- How do the MDR/IVDR and AIA mandate continuous performance monitoring mechanisms for MDAI?

The MDR/IVDR require manufacturers to establish post-market surveillance systems to monitor the performance and safety of medical devices and IVDs, including AI-enabled ones, after they are placed on the market (MDR, Article 83, IVDR Article 78). Similarly, the AIA mandates post market monitoring and post-market monitoring plan to actively and systematically collect, document, and analyse relevant data on the performance of high-risk MDAIs through their lifetime and ensure continuous compliance with requirements of Article 8 to 15 AIA (Article 72 AIA). The post-market monitoring system shall be proportionate to the nature of the technology and the risks of the system, to ensure ongoing compliance with safety and performance standards (Article 72 AIA).

.

- What new dimensions does the requirements of the AIA introduce to the existing post-market surveillance requirements of the MDR/IVDR?

Firstly, what doesn’t change is the obligation of the manufacturers for post-market monitoring of MDAI. Both MDR and AIA include obligations on the manufacturer to establish and document a post-market monitoring system that is risk-based and clearly established in the QMS that the device continues to operate as intended.

The major monitoring change required by the AIA will be the need, where relevant, to detect for interaction with other AI systems, including other devices and software. Furthermore, deployers in accordance with Article 26 are required to monitor the operation and where relevant inform manufacturer in accordance with Article 72 on post-market monitoring. Additionally, the definition of a serious incident in accordance with Article 3(49) AIA covers “serious harm to property or the environment” which is not necessarily within the scope for similar reporting under MDR/IVDR.

In accordance with Article 72(3) AIA, the Commission shall adopt implementing act laying down detailed provisions establishing a template for the post-market monitoring plan and the list of elements to be included in the plan by 2 February 2026. In accordance with Article 72 (4) paragraph 4, the necessary elements of the post-market monitoring plan and the elements of the template adopted under the AIA may be integrated into already existing post-monitoring plan under the MDR/IVDR, provided it achieves an equivalent level of protection.

# Other questions

- Do high-risk AI systems intended solely for in-house use require the CE mark?

A health institution developing an AI system for in-house use only, in accordance with Article 5(5) MDR on in-house manufacturing is not required to undergo a third-party conformity assessment and shall not CE mark the AI system. Therefore, third party conformity assessment condition for classification as high-risk AI system under Article 6(1) is not fulfilled. To ensure compliance of MDAI for in-house use only, the following conditions, in accordance with Article 5(5) MDR must be met:

- Manufactured and used within a health institution:  The system must be manufactured and used within the health institution. Its use in other facilities that do not belong to the institution is not permitted.

- Specific patient population: The system must be designed to treat a specific population of patients.

- Lack of commercial alternatives: The system should only be developed if no equivalent commercial alternatives are available.

By adhering to these requirements, health institutions can ensure the appropriate use of in-house developed AI systems while maintaining alignment with regulatory expectations.

Still, such company or health institution, where applicable, must meet the other AIA obligations, e.g. on prohibited practices or transparency.

Should ‘in house’ MDAI manufactured and used only within health institutions be classified as a high-risk AI system? As stated in question 2, one of the conditions to determine if an AI-system is high-risk establishes that the MDAI must be subject to a third-party conformity assessment by a notified body designated under the MDR/IVDR. As a consequence of EU MDR/IVDR Article 5(5), medical devices manufactured and used only within health institutions (in-house developed medical devices) established in the Union are not subject to third-party conformity assessment, provided that the conditions of Article 5(5) are met. Therefore, such a MDAI is not classified as a high-risk AI system. Nevertheless, other AIA obligations apply including but not limited to prohibited practices.

- Is the AIA applicable to custom-made devices CMD?

The AIA applies to all medical devices that fulfils the definition of an AI system, even if those devices do not require the CE mark or third-party assessment. If the aforementioned medical device is custom made, the AIA therefore still applies but would not be considered high-risk (application of Article 50 still applies). However, if the custom-made is class III under the MDR, it is required to undergo third party assessment, then it would be considered high-risk under the AIA as per Article 6(1).

- Are MDAI manufacturers required to define some minimum AI training e.g. to medical physics experts and radiologists to be able to understand potential risks of diagnostic support tools?

MDR/IVDR and AIA require manufacturers to ensure training of deployers using MDAI when appropriate as part of their risk management to ensure appropriate use, reduce foreseeable misuse and oversight during its deployment. One of the essential requirements of the AIA for high-risk AI systems includes Article 13 on transparency and provision of information for deployers. Among other things, as explained in recital 72 AIA transparency, including the accompanying instructions for use, should assist deployers in the use of the system and support informed decision making by them. Deployers, should be in a better position to make the correct choice of the system that they intend to use in light of the obligations applicable to them, be educated about the intended and precluded uses, and use the AI system correctly and as appropriate.

In addition, when oversight measures are identified commensurate with the risk, level of autonomy and context of use, the natural persons to whom oversight is assigned are enabled to understand the capabilities and limitations and be able to duly monitor. Manufacturers should advise on education and training which would provide a sufficient understanding on the MDAI interpretability of the generated output to mitigate reasonably foreseeable misuses. Similarly, MDR and IVDR requires manufacturers to supply information on special training required to use the device (Annex I).

Furthermore, Article 4 AIA, as explained in recital 20. obligates manufacturers and deployers of AI systems, to ensure, to their best extent, a sufficient level of AI literacy of their staff and other persons dealing with the operation and use of AI systems on their behalf, taking into account their technical knowledge, experience, education and training and the context the AI systems are to be used in, and considering the persons or groups of persons on whom the AI systems are to be used.

# Annex I – Qualification decision tree

Regulation (EU) 2017/745 of the European Parliament and of the Council of 5 April 2017 on medical devices

Regulation (EU) 2017/746 of the European Parliament and of the Council of 5 April 2017 on in vitro diagnostic medical devices

Regulation (EU) 2024/1689 of the European Parliament and of the Council of 13 June 2024 laying down harmonised rules on artificial intelligence

On the basis of the New Legislative Framework, as clarified in the Commission notice ‘The “Blue Guide” on the implementation of EU product rules 2022’,  the general rule is that more than one legal act of Union harmonisation legislation, such as the MDR and the IVDR and the AIA , may be applicable to one product, since the making available or putting into service can take place only when the product complies with all applicable Union harmonisation legislation.

MDCG 2019-11: Guidance on Software Qualification and Classification.

Commission Guidelines on the definition of an artificial intelligence system

established by Regulation (EU) 2024/1689 (AIA).

Commission Guidelines on the definition of an artificial intelligence system established by Regulation (EU) 2024/1689 (AI Act), C(2025) 924 final.

AIA Art 3(14) defines safety component as ‘a component of a product or of an AI system which fulfils a safety function for that product or AI system, or the failure or malfunctioning of which endangers the health and safety of persons or property’.

As regards AI systems that are safety components of products, or which are themselves products, falling within the scope of [MDR or IVDR], it is appropriate to classify them as high-risk under the AIA if the product concerned undergoes the conformity assessment procedure with a third-party conformity assessment body pursuant to the [MDR or IVDR].

Commission Guidelines on prohibited artificial intelligence practices established by

Regulation (EU) 2024/1689 (AI Act), C(2025) 884 final.